## Slide 1
Transmission Electron Microscopy

## Slide 2: SEM & TEM
Scanning electron microscope
Electron gun
Specimen
Electron gun
Condenser lens
Objective lens
Specimen
Projector lens
Fluorescent screen
Electron gun
Digital camera
Condenser lens
Objective lens
Specimen
Projector lens
Fluorescent screen
Electron gun
Digital camera
Transmission electron microscope

## Slide 6
Accelerating voltage
Electrons emitted by the filament are accelerated by a series of anodes. For transmission electron microscopy, accelerating voltages typically range from 60 kV to 300 kV. Higher accelerating voltages give higher resolution, but less contrast. High accelerating voltages can also result in greater specimen damage. For these reasons, studies of biological samples tend to employ low accelerating voltages (60 kV to 100 kV), while studies of inorganic materials, which often require higher resolution, usually employ an accelerating voltage of 200 kV or more.

## Slide 8
Electromagnetic lenses
All modern TEMs use electromagnetic lenses. These consist of a coil of copper wires inside iron pole pieces.
A current through the coils creates a magnetic field, symbolized by red lines in the diagram on the left. Electrons close to the centre are less strongly deflected than those passing through the lens far from the axis.

## Slide 9
Electrons emitted by the gun are accelerated, typically by 80 - 200 kV.
The condenser lenses focus the beam onto the specimen, which is introduced into the microscope through an airlock.
The objective lens forms a focussed image, which is enlarged by the projector lenses.
The image is viewed on a fluorescent screen or a computer monitor.
Components of the TEM

## Slide 10
Scheme of beam within a TEM

## Slide 11
Chromatic aberration
Chromatic aberration is caused by a lens having a different refractive index for different electron energies. It is present in all electron lenses, but can be reduced by minimising the energy spread of the electron source. Field emission sources have the lowest energy spread. There is currently interest in developing chromatic aberration correctors for TEM, but this is still at quite an early stage.

## Slide 12
Spherical aberration
A lens suffers from spherical aberration if it focuses rays more tightly if they enter it far from the optic axis than if they enter closer to the axis. It therefore does not produce a perfect focal point. This is illustrated in diagram A on the right.

## Slide 13
Spherical aberration correctors
In recent years, spherical aberration correctors have been developed for transmission electron microscopes, and this has greatly improved their resolving power. Aberration corrected TEMs can have a resolution of 0.1 nm or better.
Using these new ultra-high resolution TEMs it is possible to image the atomic structure of crystals such as graphite, as shown on the right.

## Slide 17: Transmission electron microscopy
Specimens for TEM
3 mm TEM grid
Carbon films

## Slide 22
Absorption contrast:
Heavy atoms absorb more electrons than light atoms.
Contrast in the TEM – Scattering
Biological and “soft matter”
Scattering contrast:
Some electrons blocked by objective aperture.

## Slide 23: Positive staining
Electron-dense heavy metal salt selectively stains certain features within the sample.
Block copolymer containing polystyrene and polymethylmethacrylate stained with ruthenium tetroxide.

## Slide 24: Biological TEM - negative staining
Surrounding biomolecules with thin amorphous layer of heavy metal salt. Reveals structure, and reduces structural flattening that occurs in the absence of stain.
Commonly used stains: uranyl acetate, uranyl formate, sodium/potassium phosphotungstate, ammonium molybdate.
Adenovirus negatively stained with uranyl acetate

## Slide 25
Biological TEM – preparing specimens from tissue
Microtome
Fix  dehydrate  set in resin  section  deposit on grids  stain

## Slide 26
Biological TEM – thin section
Kidney

## Slide 27
Contrast in the TEM - Diffraction
Cr23C6 - F cubic
a = 10.659 Å
Ni2AlTi - P cubic
a = 2.92 Å
Diffraction pattern from single crystal

## Slide 28
Polycrystalline materials
polycrystalline BaTiO3 spotty Debye rings
Contrast in the TEM - Diffraction

## Slide 29
Diffraction contrast
Contrast in the TEM - 2

## Slide 30: Bright Field Imaging
If the main portion of the near-forward scattered beam is used to form the image
  transmitted beam
  000 beam
  zero-order beam

## Slide 31: Transmission electron microscopy
TEM in physical science – bright field imaging
Dislocations in stainless steel foil
Platinum/alumina catalyst

## Slide 32
Ray Paths in the TEM

## Slide 33
Electron Diffraction
TEM Imaging Modes: Diffraction vs BF

## Slide 34: Dark Field Imaging
If the transmitted beam is excluded from the image formation process
  off-axis imaging
  tilted beam imaging

## Slide 35
TEM Imaging:
Ray Paths

## Slide 36
Phase contrast: lattice imaging
Contrast in the TEM - 3

## Slide 39
High Angular Annular Dark Field

## Slide 41
Energy Dispersed Spectroscopy

## Slide 46
TiO2 polymorphs

## Slide 47
TiO2 polymorphs

## Slide 48
TiO2 polymorphs

## Slide 49
MWCNT/Pd@TiO2 hybrid materials

## Slide 50
MWCNT/Pd@TiO2 hybrid materials

## Slide 51
MWCNT/Pd@TiO2 hybrid materials

## Slide 52
MWCNT/Pd@TiO2 hybrid materials
EDX mapping

## Slide 53
Fe@CNT/Pd@TiO2
hybrid materials
A) low magnification and B) high magnification HRTEM and C) HAADF-STEM images of sample Fe@CNTs/Pd@TiO2-magn thermally treated (inset of C: magnification showing the individual Pd NPs); D) EDX mapping showing carbon (red) and iron (green); E) EDX mapping showing titanium (blue), oxygen (red) and iron (green). (For interpretation of the references to colour in this figure legend, the reader is referred to the web version of this article.)

## Slide 54
Pd@CeO2 HIERARCHICAL STRUCTURES
Montini et al. J. Am. Chem. Soc. 132 (2010), 1402-1409.
Hydrolysis conditions:
THF + 30 eq H2O (120 mol vs Ce) in 4 h
THF
MUA: 11-Mercapto Undecanoic Acid
Dispersible in CH2Cl2, toluene, hexane
Pd core size: 1.8 ± 0.2 nm
MUA-Pd NPs
Ce(IV) tetrakis(decyloxide)
Controlled Hydrolysis

## Slide 55
Pd@CeO2 FOR CH4 combustion
Montini et al. Science 337 (2012), 713-717.
HAADF-STEM

## Slide 56
Pd@CeO2/Al2O3 materials