## Slide 1
Deterioration of orthopedic-grade UHMWPE due to ionizing radiation
One of the most significant challenges in orthopedics has been the poor oxidation resistance of ultra-high molecular weight polyethylene (UHMWPE) concomitant with ionizing radiation sterilization. This degradation is associated with osteolysis (bone loss around the implant) due to the poor wear resistance of oxidized UHMWPE polymer used as the bearing surface in total joint replacements. Until about 1995, UHMWPE was typically sterilized with a nominal dose of 25–40 kGy of gamma radiation in the presence of air. By 1998, all of the major manufacturers had shifted their sterilization method away from gamma irradiation in air
Gamma irradiation, like other high-energy photons, generates free radicals in polymers through bond cleavage. Radicals generated in the crystalline regions of the polymer have long lifetimes, allowing them to diffuse into the amorphous regions of the polymer and undergo oxidative reactions when oxygen is available. This time-dependent free radical reaction mechanism poses serious concern for the structural degradation of UHMWPE

## Slide 2
Oxidation results in
- a decrease in molecular weight,
- an increase in crystallinity and density due to chain scission and loss of mechanical properties
Moreover, oxidation results in
- severe degradation of fatigue and fracture properties of the UHMWPE
Figure depicts an FTIR spectrum showing oxidation of an UHMWPE tibial (knee) component that has shelf-aged for 5 years following gamma radiation in air. There is a clear increase in the carbonyl ketone peak at 1720 cm−1 indicating oxidation
Evidence of most significant levels of oxidation occurred beneath the surface where contact stresses are highest. This sub-surface peak in oxidation is due to the coupled effect of free radical distribution and oxygen diffusion. The coupled effect of oxidation and contact stresses is thought to be linked to delamination wear of UHMWPE