## Slide 1
FRACTURE MECHANICS
How might implants catastrophically fail when experiencing stresses far below the yield or ultimate tensile strength? How could this outcome be affected by material behavior, manufacturing, and quality concerns?
Components fractures occur when cracks or sharp flaws cause a local extreme stress that causes material breakdown, even under relatively mild loading conditions. Flaws or cracks can either exist inherently in the material or be caused by manufacturing, which can be mitigated through inspection and quality control. Cracks can also form during loading near sharp notches designed into the component, even if the component is flaw-free.
Fracture is a failure mode distinct from the concept of strength that is commonly applied in engineering design. While strength can be thought of as a value of homogenous stress corresponding to material failure, fracture is governed by localized processes that occur near the site of flaws.

## Slide 2
Linear elastic fracture mechanics (LEFM)
Griffith derived that the actual tensile strength of a brittle material containing crack-like flaws is
where the stress to cause fracture is now also dependent on the crack size, a, and the surface energy, γ .
In a more extreme case, mussel shells are 95% microscopic calcium carbonate platelets, but the 5% of organic material and the material microstructure boosts the toughness from a value of ∼2 J/m2 for pure calcium carbonate to 1500 J/m2 . The latter is also a good example of how energy dissipation beyond surface energy contributions is important in the fracture of many materials

## Slide 3
Cracks as extreme stress concentrators
Fracture mechanics is concerned with the locally extreme stresses that occur near the tip of a crack or flaw under stress, which is typically much more severe than the stress experienced by material remote from the flaw.
Stress concentrations can occur at screw threads, geometric reductions of cross-sectional areas, rivets, holes, and notches. These stress concentrations serve as considerable design challenges for engineers as such design features are often necessary. A geometric or stress concentration factor, Kt, is used to relate the local stress at the discontinuity, σloc, to that of the stress that is applied far-field, σ∞, and is defined as follows
One geometry that can be solved numerically is the case of an embedded ellipse in an infinite plate (Inglis 1913)
where b is the half-width (half of the long axis) of the ellipse and a is the half-height (half of the short axis).

## Slide 4
The stress intensity factor, K
From the mathematical theory of elasticity it can be shown that the stress at the tip of the crack in linear elastic isotropic materials decays as a strong function of distance from the crack, r, (in polar coordinates),
where fij and gij (m) are dimensionless functions giving the angular distribution (out of the crack plane) of the stress, and A and Am are coefficients amplifying the magnitude of stress resulting from each term.
The leading term in Equation becomes singular as r approaches zero and the higher order terms are vanishingly small in comparison. Thus, in LEFM, it is typically just the first term of Equation that we use to describe the distribution of stresses in the singularity region we predicted in Equation (9.4) to exist at the tip of a sharp crack,
The coefficient K is called
the stress intensity factor.

## Slide 5
Inspecting the previous Equation one sees that the stress at a sharp crack tip (ρ ≪ a) is proportional to the global stress and the square root of crack length, a. It can be shown that K follows exactly the same dependence on global stress and crack length, i.e.,
The stress intensity factor formula in Equation is often compiled in tabular form for experimental specimens using an alternative expression,
where P is the applied load, B is the specimen thickness, and W is the effective specimen length.
The polynomial F(a/W) for the compact tension (CT) specimen

## Slide 6
Energetic methods in fracture mechanics
The energy dissipated through the creation of new crack surface makes it an attractive alternative measure of localized crack phenomena. Such an observation served as the motivation for Griffith to describe fracture mechanics through an energetic equilibrium approach
where the increment in work required to create new crack surface area dWS then must be balanced by a change in potential energy, dEP, which is supplied by internal and external work done on the body
Following the nomenclature developed by Irwin (Irwin, 1956)
dEP denotes the change in potential energy, the energy release rate, G:
which is the available internal energy that can be supplied to produce a new crack surface
Therefore
Thus, the energy release rate G is a crack growth driving force.

## Slide 7
Crack growth and resistance
The energetic treatment revealed that a certain value of crack driving force must be overcome to extend a crack under equilibrium conditions. The equilibrium condition in Equation can be rewritten as
where the material resistance to crack growth is denoted as R
However, this only describes the condition at incipient crack growth without determining whether the process is stable or unstable. If we take a derivative of the equation with respect to crack length, we can evaluate the stability of crack equilibrium. If the crack growth resistance increases at a greater rate than the driving force, we have a stable equilibrium

## Slide 8
If the stored energy in the body is sufficient to start a crack propagating, so long as it continues to exceed the crack growth resistance the crack will continue to propagate, driven solely by the stored potential energy in the body. Put another way, the energy dissipated by crack growth comes from stored energy that has already been provided prior to the onset of propagation, and the remaining potential energy in the system may exceed that required for catastrophic fracture. In this way, large bodies with small cracks can supply all the energy to sustain crack growth to failure, in a way that is not the least bit improved by their extent, even if they are stressed far below the ultimate strength.
It can be seen in the previous Figure that a material with a constant crack growth resistance immediately becomes unstable when G = R, if G is an increasing function of a. For another material with the same load applied but an increasing R curve, we see that crack growth will arrest after a stable extension of the crack. After some increasing of the load, the instability condition is met and the crack is driven to unbounded growth.

## Slide 9
The flat R curve reflects the perfectly brittle fracture assumed to occur by Griffith, where the energy consumed is derived from the surface energy of the material, and is therefore a material constant.
Increasing R curves may reflect either intrinsic or extrinsic toughening mechanisms (e.g., plastic hardening, crack closing, crack tip bridging).
The instability condition shown in Figure corresponds to a critical value of the energy release rate, denoted Gc. It is apparent that Gc may or may not be preceded by an amount of stable crack growth. However, the value of Gc may vary for a given increasing R curve, depending on the functional form of G(a), which is dependent on the geometry and boundary conditions of the loaded body.
Thus, the R curve is a material property, but Gc is in general not, though it should be noted that the R curve shape is sensitive to stress triaxiality (i.e., plane stress vs. plane strain).

## Slide 10
Superposition of G and relationship to K
G and K can be directly linked, and thus the energetic contribution from each mode of K in fact can be directly summed and then used to evaluate fracture risk. Consider the energy release rate for an unbounded center-cracked plate,
where E′ = E for plane stress and E′ = E/(1 − ν2) for plane strain. For an infinite plate one can combine with the equation
K=Y s (p a)1/2 and therefore (Y=1)

## Slide 11
Kc and Gc as intrinsic failure criteria
both K and G completely characterize the local conditions near a crack tip, and therefore both are sufficient indicators of the critical condition that results in unstable crack growth.
The critical value of K is called the fracture toughness, KIc, which is a material property and is not dependent on the global geometry. However, K is dependent on the stress triaxiality at the crack tip and is sensitive to the conditions of plane stress or plane strain.
.
The critical value of KI is typically much greater than that for KII or KIII, in that fracture under Modes II and III is uncommon, and under mixed-mode loading crack growth tends to be dominated by Mode I behavior. Thus, mixed or non-tensile modes of loading may be important in specific applications or for incipient crack growth, but it is Mode I that is typically the most critical and that therefore is used to describe the fracture failure condition.

## Slide 12
A critical sized flaw can be defined relative to the fracture toughness
Limitations of LEFM
Linear elastic fracture mechanics provides a mathematically compelling explanation for how sharp flaws in bodies can produce locally extreme stresses under load, and result in catastrophic failure of components loaded far below their material strength.
An essential assumption in this analysis was that the local breakdown of material at the crack tip does not substantially affect the local stress field and that the energy required to extend a crack in an elastic material is dominated by work required to sever bonds near the crack tip (and not due to other modes of energy dissipation).

## Slide 13
Small-scale plasticity
In LEFM it is assumed that the extent of material breakdown near the crack tip was insignificant compared to the region dominated by the stress singularity.
All materials will deform inelastically in this region, and in materials that can deform plastically, as do metals, this inelastic region may in fact engulf the singularity zone and nullify the validity of LEFM. In situations where the inelastic region (called the plastic zone hereafter) is limited in extent, LEFM is still valid, and we denote this condition small-scale plasticity (also called small-scale yielding).
Thus, the correction of the plastic zone size results in an effective increase in crack
length,

## Slide 14
This, however, should not be interpreted as challenging the validity of K under plane stress conditions. Rather, the mechanisms locally governing the critical fracture condition typically depend directly on stress triaxiality (as mentioned for the R curve above), and therefore so must the fracture toughness Kc.
The relaxation of stress triaxiality from the plane strain condition commonly results in a suppression of crack growth phenomena and an apparent improvement in fracture toughness, Kc. Figure shows how Kc drops from a peak value in plane stress to an asymptote in plane strain, where the contributions of the plane stress zones on the surface of the body are negligible.
The asymptotic value of Kc in plane strain is denoted as KIc and is also known as the plane strain fracture toughness.

## Slide 15
In order to obtain a credible measure of KIc, the dimensions of a test specimen must be substantially greater in every extent than the plastic zone size. ASTM E1820 requires that the specimen thickness (B), crack length (a), and ligament ahead of the crack tip (b) all exceed approximately 25 times the corrected plane strain plastic zone size
Specimen configuration requirements for four biomaterials
Given the specimen configuration requirements, determine the minimum specimen dimension for a fracture toughness test for each of the following materials:
(a) 9.9 mm, (b) 490 mm, (c) 32.2 mm, and (d) 0.122 mm.
The titanium specimen width is reasonable. However, the necessary specimen width for the CoCr alloy is excessive, as material billets in this size may be difficult to obtain, and force necessary to conduct the test would be extremely high. For such a material, elastic-plastic fracture mechanics methodologies may be required, as most components would likely become fully plastic prior to fracture.