## Slide 1
Regulatory affairs and testing
A medical device company is developing a material for use in vascular grafts. What should the designers keep in mind from a regulatory standpoint as they move toward developing a prototype implant?
This company will want to decide if it is planning to market the device itself, or merely supply the material to another company that will be the primary distributor of the implant.
Second, the company might want to check the FDA database to see what guidance the FDA may provide for these implants.
Third, the company will need to determine whether a device made from this material is going to be entirely novel, or whether it can be classified as substantially equivalent to currently approved devices.

## Slide 2
Total hip arthroplasty
This ball and socket mechanism results in high conformity and the articulating surface of the hip replacement experiences relatively low contact stresses (2–5 MPa). Nevertheless, a large contact area and cross-shearing motion exist in the hip joint.
A mechanism of damage associated with the wear process is the generation of a plasticity-induced damage layer at the articulating surface which leads to lamellar alignment and loss of structural integrity as the implant undergoes a cross-shear articu- lation

## Slide 3
In order to overcome the clinical challenge of osteolysis associated with wear of the bearing surface, the orthopedics community has moved toward highly crosslinked forms of UHMWPE that offer exceptional wear resistance. Crosslinking is achieved with radiation and subsequent thermal treatments that annihilate free radicals that could otherwise render the polymer susceptible to oxidation.
A challenge for the orthopedic manufacturers, however, is that the crosslinking comes at the expense of fatigue and fracture resistance

## Slide 4
Total knee arthroplasty
The primary movement of the knee joint is flexion and extension in the sagittal plane with a small amount of rotation. These motions induce a small area of contact between the femoral condyles and the tibial plateaus due to the constraint provided by the ligaments. The knee achieves much of its stability from the large ligaments that attach the femur and tibia and fibula.
The motion in the knee is also more complex than the hip, including rolling, sliding, and some rotation due to flexion and extension of the joint.

## Slide 5
The primary challenge is the propensity for wear at the articulating surface of the primary bearing pair owing to enhanced conformity and also at the backside interface between UHMWPE and its underlying alloy in the rotating bearing. As with other implant designs, increased levels of particulate debris are accompanied by a greater likelihood for bone loss around the implant due to osteolysis.
Owing to the less conforming nature of the knee anatomy, the device is expected to sustain large cyclic contact stresses that may span 10 MPa in tension to 40 MPa in compression as the knee undergoes flexion and extension
A source of failure was subsurface fatigue fracture (delamination) of the UHMWPE caused by high contact stresses and in some instances, oxidation of the polymer owing to poor sterilization methodology

## Slide 6
Historically material enhancements have often resulted in poor clinical performance. One such example was the development of carbon fiber-reinforced UHMWPE (Poly II). This material was marketed to offer improved mechanical performance over conventional UHMWPE with the assumption that increasing the strength of the bearing material would improve the life of the implant. In laboratory studies, the composite offered a 25% increase in monotonic compressive strength and an order of magnitude improvement in wear resistance over the conventional UHMWPE;
Regrettably, Poly II fared poorly in clinical practice owing to poor interface strength (adhesion) between the fiber and the matrix as well as enhanced contact stresses associated with the increased elastic modulus of the composite structure. Figure
shows a typical fracture surface with fiber debonding in the UHMWPE matrix.

## Slide 7
Fracture fixation
CoCr alloys,
stainless steels,
and titanium alloys are utilized in fracture fixation

## Slide 8
Pins are used to fasten or attach fragmented bones together and are commonly used in fractures that are too small to be fixed with screws. Pins may be removed or left in place after the fracture heals.
Wires are generally employed like metal sutures to thread or tie fractured bones back together.
Plates serve as internal splints and are screwed in place to hold the fractured bone together. These devices may be removed after healing or left in place, depending upon the nature of the fracture.
Screws are the most commonly used implants for internal fixation of fractured bone and these may be used independently or in combination with the other fixation implants. There are two basic types of screws available. Cortical screws are designed for compact diaphyseal (cortical) bone, whereas cancellous screws are designed for the metaphyseal (trabecular) bone. The former typically have a larger thread diameter and pitch and a greater difference between major and minor (shaft) diameters in comparison to cortical screws, providing more surface area for bone attachment. These screws are intended for use in metaphyseal fixation, where bone is softer.

## Slide 9
Cannulated screw sizes vary from 4.0–7.5 mm and cancellous screws are typically between 4.0–6.5 mm.
Cannulated screws have a hollow central shaft. The main advantage of cannulated screws is that they can be inserted over a guide wire or guide pin. The diameter of the guide pin is much smaller than the cannulated screw and can be more accurately placed using fluoroscopy in the operating room. In addition, given its small diameter, the guide pin can be reinserted several times if necessary for accurate placement without excessive damage to bone.
Rods are sometimes referred to as intramedullary rods (or nails) and are as a rule employed in the repair of long bones; they are typically inserted through the center of the long axis of the bone. Like the other implants, the rods may be left in place once the fracture is repaired.

## Slide 10
Spinal implants
There are two primary categories of medical devices used in the spine, and this includes spinal fixation and disk replacement.
The motivation of using an implant to replace an intervetebral disk is to remove the damaged disk while preserving motion in the spine
However, the long-term clinical success of spinal arthroplasty remains uncertain.
An anatomical disk replacement device seems like a logical design; however, in order for the rubber core to be mounted it was necessary to vulcanize it to the titanium plates. It was at the titanium endplate-rubber core interface where failures occurred owing to repetitive shear forces induced by physiological loading and kinematics of the spine. An additional complication in the rubber implant was the generation of particulate debris that resulted in an inflammatory response and propensity for osteolysis in patients

## Slide 11
After a short time on the market, AcroFlexTM was recalled and the implant was revised using a silicone elastomer in place of the rubber core. The modified device received much criticism for its use of silicone at a time when this material was under scrutiny during the Dow Corning litigation involving silicone breast implants, and the revised AcroFlexTM device was never implemented in clinical practice. While this implant was not clinically successful, it provided the spinal implant community with a foundation of knowledge that would lead to development of disk replacements with multiple components that allow for motion without large deformation of the implant.
Spinal implant designs
The Charite intervertebral disk is the most widely used lumbar interbody implant currently available.
The design utilizes a biconvex UHMWPE insert that acts as a mobile bearing within endplates made of a cobalt-chromium alloy.

## Slide 12
The Charite intervertebral disk
The Maverick intervertebral disk
implant utilizes CoCr alloy articulation with a fixed posterior center of rotation to better mimic the kinematics of the intervertebral motion segment and larger endplates to enhance stability of the implant

## Slide 13
The Flexicore intervertebral disk
The FlexicoreTM is under evaluation in a multi-center clinical trial
Corrosion-induced fracture of a double modular hip prosthesis (design failure)
A cementless hip stem with ceramic-on-ceramic articulation implanted in a 30-year-old man with a BMI of 29
fractured after two years in vivo due to a fall on his right hip

## Slide 14
The stem and neck are both manufactured from an orthopedic Ti alloy, and are offered in a range of sizes to accommodate a range of leg length and femoral adjustments.
Stem-neck design (modularity) may contribute to susceptibility to corrosion-induced fracture.
Ti6Al4V alloys are generally protected against corrosion in vivo by a passive oxide layer on the surface; however, fretting occurred at the modular junction in the crevice of the neck-stem interface and wore away the protective oxide film. This design exacerbated the corrosion process, resulting in pitting and crack nucleation at this site. The bending stresses and likelihood for failure are greater in longer neck designs because the force scales with neck length.

## Slide 15
The progression of events leading to the fracture in this implant is illustrated in Figure
The fracture of this modular neck is a representative case of the trade-off in modular implant design: modularity enables the surgeon to better optimize the patient’s biomechanics and anatomy, but additional mechanical junctions can lead to corrosion processes, resulting in fracture. This case shows that there is a risk of fracture at the stem-neck junction when using double modular devices, particularly with long neck lengths in heavy patients

## Slide 16
Sulzer recall (manufacturing failure)
In September 2000, Sulzer Orthopedics (Austin, Texas) began receiving adverse reports from surgeons who were getting complaints of leg pain from their patients recovering from recent total hip replacements procedures
31,000 acetabular shells manufactured between 1997 and 1999, were recalled back to the company.
17,500 had already been implanted.
It was determined that a small amount of mineral oil-based lubricant had leaked into the machine coolant during the manufacturing process and the lubricant remained as residue on the shell. This retained oily residue on the defective implants inhibited the osseointegration necessary for fixating the shell in the bone of the pelvis
Only 2,400 implants were removed and revised as of 2005. Sulzer agreed to pay a $1 billion settlement in a class action lawsuit filed in 2002. Patients who had undergone a revision surgery were awarded $160,000. After the lawsuit settled, Sulzer Orthopedics changed its name to Centerpulse Orthopedics.

## Slide 17
Cardiovascular devices
Vascular grafts
Vascular grafts are needed when a patient is suffering from severe atherosclerosis and whose effects cannot be relieved through balloon angioplasty or stenting. This condition is known as angina and the resulting surgery is called coronary artery bypass graft (CABG) surgery. A person can have multiple bypass grafts in a single surgery.
Grafts are also used when an aneurysm occurs. An aneurysm is a bulge in a blood vessel due to thinning and weakening of the vessel wall. They commonly occur in the aorta and at the base of the brain. In the aorta, an endovascular stent-graft can also be implanted to serve as an alternative conduit for blood flow.

## Slide 18
Current available vascular grafts can be divided into natural and synthetic options.
Natural options include the saphenous vein and the umbilical vein, which is sometimes enclosed in a polyester mesh to increase wall strength. These grafts show good durability and a lack of surgical complications. Unfortunately, 20–30% of patients requiring a graft do not have a suitable saphenous vein for the transplantation process
This fact, combined with a need for larger diameter replacements such as for the aorta, drives the development of synthetic grafts.
These grafts, made of Dacron or expanded PTFE (ePTFE), can be knitted, crimped, or otherwise processed to maintain their structure. ePTFE has the necessary mechanical strength to perform well in this application and has been shown to have low occlusion due to thrombosis).
PET fabrics have high stiffness and low extensibility. Woven fabrics have lower porosity than knitted and do not require pre-clotting, while knitted grafts with velour are believed to have better tissue integration and to enhance endothelialization.

## Slide 19
knitted
woven

## Slide 20
Thrombosis occurs readily in synthetic grafts of diameters less than 6 mm, and so these are not generally implanted below the knee or in CABG surgery

## Slide 21
There are only two total artificial hearts approved for use by the FDA. The Syncardia temporary Total Artificial Heart (Syncardia Systems, Inc., Tucson, Arizona) was approved in 2004 and is used as a bridge-to-transplant. The AbioCor Replacement Heart (AbioMed, Inc., Danvers, Massachusetts) was approved under a Humanitarian Device Exemption in 2006 and is used only in patients with severe end-stage heart disease who are not eligible for a heart transplant. For more information on TAHs, Blitz and Fang (2010) provide a strong overview. The Syncardia TAH is shown in Figure 14.16.
TOTAL ARTIFICIAL HEARTS

## Slide 22
Nitinol stent strut failure
(structural failure)
In current practice, it is common to use flat plane X-ray detection or fluoroscopy, a technique that uses X-rays to create a continuous, real-time image. Other techniques, such as angiography or CT angiography, are further variants of X-ray and fluoroscopy
It was found stent fracture in 65.4% of nitinol stents, mainly non-coil designs. still associated with restenosis.

## Slide 23
Oral and maxillofacial devices
Dental implants
Three types of dental implants are utilized in modern dentistry
– subperiosteal, blade (plate form) endosteal, and root-form endosteal. All three types of implants are composed of two main parts – the implant body, which is the portion of the implant placed within or atop the alveolar bone, and the abutment(s), which attach to the implant body and are transmucosal, extending through the gingiva.

## Slide 24
Endosteal implants are implants in which the implant body is embedded within bone. In blade-form endosteal implants, the implant body is a thin, blade-like metal framework that typically forms a single piece with the abutments. Blade-form implants may be used in healthy or mild to moderately atrophied alveolar
On the other hand, the implant body in modern root-form implants is either a cylinder or screw that resembles the natural tooth root. In screw-type implants, the implant body is further comprised of an apex, body, and a crest module

## Slide 25
Masticatory loads, which are concentrated in the posterior region, are primarily compressive and applied along the long axis of the tooth, orthogonal to the occlusal plane. These forces vary with type of food, and, depending upon the device used to measure them, typically range between 100–250 N in the molars, 40–150 N in the premolars, and less than 50 N in the incisors. The maximum bite force (MBF), measured as the maximum force generated during direct contact between the maxillary and mandibular dentition, is significantly greater. Like masticatory forces, the MBF is highest in the posterior molar region and decreases progressively the further anterior the teeth. However, it is much more variable than masticatory forces, and has been measured in the range of 300–1000 N for the molars, with only 30% of those loads observed in the incisal region. Supra-physiologic loads as high as the MBF occur rarely, and account for negligible percentage of the loading experienced by teeth, both natural and artificial.

## Slide 26
FDA recall of Proplast-Teflon temporomandibular joint replacements (regulatory and materials failure)
The massive failure of the Vitek-Kent TMJ prosthesis in the mid-1980s has been dubbed by many as one the greatest American medical disasters in history
For many decades, controversy surrounded the question of whether the TMJ was a weight-bearing joint
In light of its low coefficient of friction, there was a growing interest in using Teflon, Dupont’s trademark for polytetrafluoroethyelene (PTFE), in biomedical applications in the 1960s.
Dupont issued a warning to device manufacturers in 1968, stating that its Teflon products possessed poor wear properties, and, in 1966, Charnley had written a letter emphasizing the poor performance of Teflon in total hip replacements

## Slide 27
Despite these warnings, Vitek pressed on with its Proplast-Teflon material, taking the position that the TMJ was either unloaded or only minimally so, even under degenerative conditions
Whereas clinicians and Vitek failed in their scientific responsibilities, the FDA did not uphold in its regulatory duties. Though Silastic and Proplast were both polymers, the materials possessed very distinct physicochemical and mechanical properties – Silastic is a rubber, while Proplast is a semi-crystalline polymer composite. Hence, claims of substantial equivalence should have been rejected, and a full PMA should have been requested from Vitek for its TMJ prosthesis.

## Slide 28
Soft tissue replacements
Synthetic ligament
Ligaments attach bone to bone and are generally responsible for restraining motion and for transferring load
Although the geometry of a ligament seems quite simple, the search for a replacement has not yet managed to combine resistance to torsion, abrasion, and cyclic loading.
A healthy ACL will undergo 1–2 million cycles per year with an average peak load of 400 N, although the load in the ACL can be as high as 630 N during activities such as running
Research shows that the ACL can withstand a maximum load of 1716–2160 N and has a stiffness of 182–322 N/m
Synthetic ligaments are Class III devices

## Slide 29
Early failed replacements include a PTFE/carbon composite called Proplast (Vitex Inc., Houston, Texas) and the PP Polyflex (Richard, Memphis, Tennessee), both of which were withdrawn from the market due to rupture problems and inflammatory reactions.
In 2004, there were 200,000 ACL reconstructive surgeries in the United States. The traditional method for replacement is to use a donor graft from another location, such as the patellar tendon. This technique has certain drawbacks, including limited graft size, morbidity at the donor site, and a long rehabilitation period for the patient after surgery

## Slide 30
Sterilization of allografts, commonly achieved through gamma irradiation or EtO gas exposure, can change their mechanical properties, and these implants are known to weaken 2 weeks to 4 months after implantation
Permanent replacements require high-strength, highly fatigue-resistant materials. Two examples of these are the Gore-Tex Cruciate Ligament Prosthesis and Styker-Dacron implants. The Gore-Tex replacement was made of a single PTFE fiber arranged in loops. It had a failure load of 4800–5300 N and a stiffness of 322 N/m. It was introduced in the early 1980s and approved by the FDA in 1986; however, it was withdrawn in 1993 due to problems with rupture and local inflammatory reaction (FDA, 1993). The Stryker-Dacron prosthesis was a woven PET cable with a velour tube covering. It had a failure load of 3631 N and a stiffness of 420 N/m. It was approved by the FDA in 1988 but withdrawn in 1994, with clinical data from one study showing a 40% rupture rate after 18 months of implantation

## Slide 31
Some devices are characterized as permanent augmentation devices, meaning that they provide support for the existing ACL while it heals, and in case it does not recover full strength. Examples of these include the Kennedy Ligament Augmentation Device (LAD) and the Trevira.
The Kennedy LAD is an 8 mm diameter ribbon of woven PP, designed to allow for progressive load transfer from the implant to the healed ligament. It has a failure load of 1500–1730 N and a stiffness of 280 N/m. It was approved by the FDA in 1987. The Trevira prosthesis is. made of woven PET, it is twisted to achieve a specific stress-strain relationship. The untwisted failure strength is 4000 N, but the twisting reduces this to 1866 N. The stiffness is 68 N/m. It has been implanted since 1980