## Slide 1: Physisorption, surface area and porosity

## Slide 2
Brunauer Emmett and Teller (BET equation):
n amount of adsorbate (g) at the pressure P
nm amount of adsorbate corresponding to the formation of a monolayer
C BET constant
P0 saturation pressure
Mw è Molecular weight of the adsorbate Na Avogadro Number am “cross section area” of the adsorbate.

## Slide 5
What is a Gas Sorption Analyzer?
Does it actually measure surface area and pore size?
NO!! It simply records various pressures of gas in the sample cell due to adsorption and desorption. The instrument then calculates the amount (as STP volume) of gas adsorbed/desorbed.
Surface area, pore size are calculated by PC software
Pressure measurements are critical!

## Slide 6
Vapour pressure at 77 K
Kr = 267 Pa N2 = 101325 Pa
The gas adsorption amount in the volumetric method is calculated from the difference between the number of dosed molecules and number of unadsorbed gas molecules at equilibrium pressure.
In other words, by assuming both molecular sizes equal and adsorption pressure equal to 50 Pa, it is necessary to measure the pressure change of 0.16% (=50/30400) for N2 and 63% (=50/80) for Kr at the same relative pressure (P/Po=0.3, N2=30400Pa, Kr = 80Pa).
Obviously, it is easier to measure larger pressure change, and therefore, the accuracy of the measurement is better. For this reason, the lower the saturation vapour pressure at the adsorption measurement temperature, the more accurate the measurement of low surface areas and so Xe, with vapour pressure equal to 0.23 Pa at 77 K, is still a better gas to measure low surface areas. Unfortunately, Xe is much more expensive than Kr.
Low Surface Area Materials
Nitrogen 0.162 nm2
Krypton 0.210 nm2
https://www.microtrac.it/it/applicazioni/conoscenza-base/method-of-small-surface/

## Slide 7
https://andyjconnelly.wordpress.com/2017/03/13/bet-surface-area/

## Slide 8
Degas
Important step before measurement of surface area or pore size/volume
Surfaces armust be ‘cleaned’ of water/organic vapours in two ways:
With heating under a vacuum
Under a flow of dry, inert gas.
No modification must be induced

## Slide 9
What is the result?
The values on the y-axis are calculated from pressure measurements (and temperature values)
Desorption curve may overlay on,
or appear to left of,
the adsorption curve

## Slide 10
What to Analize?

## Slide 11
What to Analize?

## Slide 12
What to Analize?

## Slide 13
What to Analize?

## Slide 15
Total Pore Volume

## Slide 16
Pore Size: Kelvin Equation

## Slide 17
Pore Size: Kelvin Equation

## Slide 19
ADSORPTION PLOT for Ce0.5Zr0.5O2

## Slide 20: Chemisorption, surface area and metal dispersion

## Slide 21
Chemisorption for probing metal surface areas
Objective
determine number of accessible metal sites
Exploit high reactivity of metal surface atoms
- let them react with small molecules (H2, O2, CO, etc.)
- count the number of molecules that have reacted
- knowing the stoichiometry of the reaction you obtain the number of metal surface atoms
Choice of the probe molecule
- must be selective for reaction with metal vs. support
- must be selective for reaction with surface atoms (not with bulk)
- reaction must be «irreversible»

## Slide 25
n = number of surface metal atoms that are occupied by the chemisorption of a gas molecule.
ex : H2 on Pt : n = 2 because of dissociative adsorption: 1 molecule H2 dissociates to 2 H atoms adsorbed on 2 Pt atoms
v adsorbed gas volume
Vm volume of a mole of metal atom
Na Avogadro’s number
S surface of one metal atom
Metal surface atoms (N(s)M) and metal surface area (SM)

## Slide 26
Chemisorption stoichiometry
Nanoparticle < 1nm  n < 2
corner atoms are usually more active than atoms on the flat surfaces
stoichiometry of chemisorption may also change

## Slide 27
Rh/Al2O3

## Slide 28
Chemisorption and Physisortion: the role of support in nanocomposites
Pressure
Gas adsorption
Physisorption is investigated at 77 K but occurs at higher temperature

## Slide 29
Chemisorption and Physisortion: the role of support in nanocomposites
Pressure
Adsorption

## Slide 30: 1.1% Pt/g-Al2O3
Adsorbimento a 573 K di H2. Curva (a) adsorbimento su 2.031 g di catalizzatore, (c) adsorbimento su 1.00 g di Al2O3, (d) adsorbimento su quantità di supporto corrispondente a quella del campione (2.00 g) ed espressa nella stessa scala della curva (a), (b) differenza tra la curva a e d.