## Slide 1: UV-VIS spectroscopy

## Slide 2: UV-Vis spectroscopy
visible（400 – 800 nm)
ultraviolet（< 400 nm）-- Near-UV（300 - 400 nm)
Deep-UV（ 200 - 300 nm)
Far-UV（< 200nm）
Electronic transitions are observed
Ozone absorption
Oxygen absorption
Doesn't arrive from space
Does not propagate in air
小川技研 3.1光波長と色 （2017.1.23改定）
http://wwwb.pikara.ne.jp/ogawa-giken/image_process/image_031.html
400
300
200
O3
O2
FUV
near-UV
DUV
energy
VIS
800
25000
33000
50000
12500
Wavelength
/nm
Wavenumber
/cm-1

## Slide 3: What can we learn from UV-VIS spectroscopy
Quantitative analysis
  Beer-Lambert law Sensitivity of about 10-6 mol L-1
  Simultaneous multi-component quantification
Understand the electronic states
  HOMO-LUMO gap
  Molecular orbitals
  → Selection rule of chemical reactions = Woodward-Hoffman rule
Frontier orbit theory
Can identify molecules
  not good at investigating the details of structures
  and functional groups compared to infrared spectra.
  Contains information about intermolecular interactions
  and associated structures
  J-aggregates/H-aggregates
H-aggregate
Shorter wavelength
than monomer
J-aggregate
Longer wavelength
than monomer

## Slide 4: Selection rule for observing electronic transitions
Bohr‘s frequency condition
The wavelength of light with energy matching the energy difference (transition energy) between the ground state and excited state of the molecule is strongly absorbed
The motion of electrons within molecules is quantized. This is because the quantum properties of the atomic nucleus and electrons are prominent in tiny spaces such as molecules, and the orbits that the electrons can take are limited. The energy of an electronic state is determined by what kind of spin an electron exists in an orbit with what kind of energy.
DE＝hn
=Energy of light
Energy interval
Molecular states

## Slide 5: Selection rule for observing electronic transitions
Light is absorbed when the excitation dipole moment is not zero.
The probability that an electronic transition occurs is proportional to the transition dipole moment . For simplicity, if we assume that the electric field of light oscillates in the spatial coordinate z direction, the transition dipole moment
the z component of the dipole moment of the molecule
(zi is the z component of the i-th electronic coordinate)
and the electronic wave function before and after the transition, (shows the complex conjugate of .
Notes: In other words, whether or not an electronic transition occurs is determined by the electron configuration within the molecule, and if the transition dipole moment is 0, no light will be absorbed and no electronic transition will occur even if Bohr's conditions are muched.

## Slide 6: Absorption spectra
wavelength
Wavenumber
Transmittance（％）
100
No abosorption
If there is absorption signal downwards
wavelength
Wavenumber
Absorbance
0
No absorption
If there is absorption signal upwards
Incident light I0
Transmittance light
I

## Slide 7: Beer-Lambert law
Lambert’s law
Beer’s law
Beer-Lambert law
Absorbance is proportional to the optical path lengt, x.
;molar absorption coefficient
Absorbance is proportional to the volume molarity, C.
molarity
/mol L-1
A:C,dx
dIx= kIxCdxI
Absorbance is proportional to
optical path length x and concentration C of the sample

## Slide 8: Quantitative analysis using UV-Vis spectra
Beer-Lambert Law
Molarity can be determined by measuring absorbance
The measured substance must be uniform and
their molar absorption coefficient does not change within the measurement concentration rangeT
Condition

## Slide 9: Error estimation of absorbance ーTwyman-Lothian’s plot
Relative error of transmittance
Relative error of Absorbance
⇦relative error of observed intensity
Values that can be
measured reliably
Even in absorbance, the same relative error occurs only between 0.1 and 1.0.
Absorbance A
Relative error of Absorbance
Relative error in absorbance when relative error in transmittance is 0.001.

## Slide 10: Simultaneous quantification of multiple components
When the maximum wavelengths are far apart and
there is no interaction between the components
condition
Even in a solution where a few components are mixed, all the components can be quantified simultaneously from the absorbance at the corresponding number of wavelengths of the spectrum.
For the case of 2 components
Simultaneous equations with two wavelengths
Solving for concentration
Spectrum of
Mixed solution
Spectrum of Y
Spectrum of X
wavelength
Absorbance
Additivity of Beer-Lambert laws
x

## Slide 11: instruments
A UV-visible spectrometer is divided into
light source,
spectrometer,
sample section
detector.
Light
source
slit
spectrometer
slit
Sample section
detector
detector
spectrometer
Sample section
Light
source
monochromator
polychromator

## Slide 12: Light source
tungsten lamp（350-2600 nm）
D2 lamp （190~400 nm）
others
Xenon flash lamp
discontinuous/high intensity （190~2000nm）
Super continuum (SC) light source:
coherent broadband light, short pulse
(very expensive) (400~2400nm)

## Slide 13: Spectrometer
Separates light into wavelengths
with a spectroscopic element
Utilizes wavelength dispersion of refractive index.
light of different wavelengths
refract to different angles
α
β
Light
source
slit
spectrometer
slit
Sample section
detector
detector
spectrometer
Sample section
Light
source
monochromator
polychromator
Using interference of diffracted light
Light of different wavelengths strengthens each other at different angles of reflection
prism
grating

## Slide 14: Sample compartment
Single beam
① Measurement without sample I0
② Measurement with sample I
Double beam
① Measuring the difference in detecting sensitivity between two optical paths
② Simultaneous measurement
Measurement without sample I0
Measurement with sample I
Example of double beam set up
With sample
Without sample
=reference
Spectrometer

## Slide 15: Detector
Si photodiode (1100-180 nm)
  Sensitivity is stable over a wide range of wavelengths (sensitivity is lower than PMT)
Photomultiplier tube（PMT) (200-900 nm)
High sensitivity
Photodiode for Nera-Infrared light
  InGaAs (800-1700 nm)
  PbS (1000-3500 nm)
multichannel detector for polychromator
・Charge-Coupled Device (CCD) (800-400nm)

## Slide 16: Do you remember?
Which of the following sentences regarding the Lambert-Beer law is incorrect?
  Beer-Lambert law states that transmittance is proportional to the concentration of solute and the optical path length.
  To apply this law, the object to be measured must be uniform.
  Under certain conditions, it is possible to quantify the concentrations of the three components using this law.

## Slide 17: Checking
Which of the following sentences regarding the Lambert-Beer law is incorrect?
  Beer-Lambert law states that transmittance is proportional to the concentration of solute and the optical path length.
  To apply this law, the object to be measured must be uniform.
  Under certain conditions, it is possible to quantify the concentrations of the three components using this law.
absorbance

## Slide 18: Materials used for measurement - spectroscopic cell

## Slide 19: Materials used for measurement – Selection of solven
1) Can dissolve the measurement sample
It is enough to dissolve about 10-2 mol L-1 (in some cases, even less is OK)
2) No absorption in the measurement wavelength range
3) No reaction occurs
4) If volatile
Pay attention to changes
in concentration
5) Polar/nonpolar solvents

## Slide 20: Preparation of solution -key process for calibration
Solvent selection (previous slide)
Precautions during preparation
  Pay attention to accuracy when weighting - take at least 0.2g if possible
Dilution is more accurate than small weighting when done
with a volumetric pipette + volumetric flask.
When diluting a solution of unknown concentration, adjust the dilution rate by diluting in 10-fold increments so that it falls within the appropriate absorbance range.
  (Absorbance 3~4 may be around 10 in practice)
Determine the amount of solution to be prepared depending on the size of the cell.
  (Approximately 4ml is required for 1cm cell)
If the accuracy of the electronic balance is 1mg

## Slide 21: Measurement (1) Setting experimental conditions
Slit width
Small slit width → High monochromaticity of light = wavelength resolution of measurement
→The amount of light used is small = it takes time to measure
Determine the appropriate slit width based on experimental purposes
Data acquisition interval/ Scan speed
Set the data acquisition interval smaller than the wavelength interval corresponding to the slit width
The scan speed determines the time it takes to process one point of data.
Relationship with slit width is important
Measurement range Various switching wavelengths
If the measurement range is wide, it will take time to measure.
If the measurement range is wide, it is necessary to switch the detector and light source.
If the measurement range is wide, an appropriate high-order optical filter is required.
Due to the characteristics of the diffraction grating, 800nm ​​primary light and 400nm secondary light reach the same location.
↑To measure this accurately, filter the short wavelength side.

## Slide 22: Measurement (2) Baseline/reference light
The most influential measurement in an experiment
The reference light should compare to that in previous measurements.
For single beam= reference beam measurement
For Double beam baseline
= correction of detecting sensitivity on two cite
Measuring pure liquid
Measure air (without anything)
Measurement of solutions
Measurement with solvent
= Removes solvent effects and cell effects
The effect of reducing the solvent in the solution appears

## Slide 23: Measurement（３）sample placement
Negative absorbance
= more light when measuring sample than reference?
Reflection by cells changes
or
Effect of decreasing solvent concentration
Scattering by the sample
= decrease light in all wavelength → baseliene of absorbance fluctuate
Eliminate scattering effects using an integrating sphere etc.
Data analysis removes increased absorbance
sample
reference
absorbance
By scattering
Affect quantitative analysis
Absorbance when there is no absorption

## Slide 24: Data analysis (1) Calibration curve
Measure absorbance at 4 to 5 points with known concentration
yi(λ）=axi + b
←Zero point deviation due to measurement
(residual error)
absorbance
Molarity/10-6 mol L-1
Unknown absobance
Calibration point
Analyzed molarity
Determine a and b by measuring known concentrations.

## Slide 25: least squares method

n
i
i
s 
2
y  (ax  b)
Determine (a, b) such that the sum of squared residuals takes the minimum value
x
y = ax + b
( xi, yi )
i
i  yi  (axi  b)
minimize the deviation of each data (xi, yi) from the approximate line.
7
Determine a and b by measuring known concentrations.
y

## Slide 26: least squares method
s is a quadratic expression of a and b
The coefficient of a2 is x2, the coefficient of b2 is 1 (both positive values)
→The minimum value is unique
2

n
 i1 
 i
n
i1
 i
n
n n
 i i  i  i
x 
x2  
n
x y  x y
n
a  i1 i1 i1
a or b
To determine (a, b) when s takes the minimum value..
Solve simultaneous equation
2
2
2

n
 i1 
 i
n
 i i1
n
n
n n
 i  i  i i  i
x 
x  
n
x y  x y x
b  i1 i1 i1 i1
・・・(1)
8
s

n
i
i
s 
2
y  (ax  b)

## Slide 27: Data analysis (2) Difference spectrum in solution
Extract the target component from the difference using the additivity of absorbance
A （solute+solvent）＝A（solvent）+ A（solution）
A（solute）＝ A （solute+solvent）ー A（solvent）
Measured spectrum
of solution
Calculated from spectrum of pure solvent
A（solute）＝ A （solute+solvent）ー A（pure solvent）
Measure the concentration of the solvent (amount of solvent) when preparing the solution in a volumetric flask.
cannot be measured
separately
Additivity of
Beer-Lambert laws
Measured spectrum
of solution

## Slide 28: Data analysis (3) derivative spectrum
derivative
additive baseline
b
linear baseline
cλ
Remove
Remove baseline shift
1st derv.
2nd derv.
Gaussian
wavelength
wavelength
wavelength
1st derv.
2nd derv.
1st derv.
2nd derv.
Notes: In the first derivative, the peak position crosses 0, so it is useful for determining precise peak positions.
In the second derivative, the peak becomes sharper than the original peak. Separation becomes possible when two peaks overlap
The height of the peak in the derivative spectrum is proportional to the height of the original peak.
However, care must be taken to reduce the height of the differential spectrum when the spectral linewidth is large.

## Slide 29: equipment（１）Integration sphere
Reference
light
Standard reflector
(standard scatterer)
Standard reflector
(standard scatterer)
detector
Sample
Sample
Standard reflector
(standard scatterer)
Reference
light
detector
Measurement for suspended solutions
Measurement for diffuse reflection
Notes: An integrating sphere is an element whose inner surface is spherical and whose inner wall is covered with a highly reflective material. The light taken into the integrating sphere is uniformly scattered by the wall surface and guided to the light-receiving section where the detector is placed. In addition to the light-receiving section, two entrance and exit holes are made in the wall surface so that the optical path of the spectrophotometer passes through. Using this device, absorption spectra of suspended solutions and diffuse reflection spectra of solid samples can be measured.

## Slide 30: 測定法（２） 光ファイバー

## Slide 31: データ解析 （２）差スペクトル
・二つのスペクトルを比べて
ｂの試料に多く含まれるものの吸収 → +
ａの試料に多く含まれるものの吸収 → －
例えば反応が進むときの時系列データ
aの試料＝反応前のスペクトル
bの試料＝反応後のスペクトル
→反応物の吸収 －
生成物の級数 +