## CTF calculation
INPUT |  |  | OUTPUT
Frequency (lp/mm) | minimum | maximum | modulation
0.05 | 0 | 1 | 1
0.1 | 0 | 1 | 1
0.15 | 0 | 1 | 1
0.25 | 0 | 1 | 1
0.35 | 0 | 1 | 1
0.5 | 0 | 1 | 1
0.7 | 0 | 1 | 1
1 | 0 | 1 | 1
1.4 | 0 | 1 | 1
2 | 0 | 1 | 1
3 | 0 | 1 | 1
4.5 | 0 | 1 | 1
6 | 0 | 1 | 1
8 | 0 | 1 | 1
10 | 0 | 1 | 1

## LSF to MTF calculation
Date: 14/02/2023 |  |  | INPUT | OUTPUT
Author: Luca Brombal
REAL SPACE |  |  |  | RECIPROCAL - FOURIER SPACE
Pixel size (mm) | Position (mm) | Pixel number | LSF | Frequency sampling (lp/mm) | Frequency (lp/mm) | FT(LSF) | MTF | auxiliary 4 plotting
0.02 | 0.02 | 0 | 0 | 0.19531 | 0 | 0.69196845309 | 1.0608 | TRUE | 0
| 0.04 | 1 | -7.1583e-05 |  | 0.19531 | -0.620402667767699+0.201455622694088i | 1 | TRUE | 0.19531
Number of input data | 0.06 | 2 | -0.00062996 | Nyquist frequency (lp/mm) | 0.39062 | 0.491167230092664-0.351232155898772i | 0.9257 | TRUE | 0.39062
256 | 0.08 | 3 | 0.00044786 | 25 | 0.58594 | -0.328773322296169+0.439889608747544i | 0.84192 | TRUE | 0.58594
| 0.1 | 4 | 7.3194e-05 |  | 0.78125 | 0.176858839367354-0.475634274077496i | 0.77795 | TRUE | 0.78125
| 0.12 | 5 | 0.00072563 |  | 0.97656 | -2.44411015305439E-002+0.471563541940732i | 0.7239 | TRUE | 0.97656
| 0.14 | 6 | -0.00052324 |  | 1.1719 | -0.105270143894249-0.442362571060159i | 0.69711 | TRUE | 1.1719
| 0.16 | 7 | -6.8588e-05 |  | 1.3672 | 0.239691480199043+0.366617529333171i | 0.67151 | TRUE | 1.3672
| 0.18 | 8 | 0.00026191 |  | 1.5625 | -0.322576769907089-0.240964966227651i | 0.61727 | TRUE | 1.5625
| 0.2 | 9 | -0.00033309 |  | 1.7578 | 0.347682175679155+0.120361615607487i | 0.56405 | TRUE | 1.7578
| 0.22 | 10 | 0.00039067 |  | 1.9531 | -0.359205252998011-1.88549115458766E-002i | 0.55144 | TRUE | 1.9531
| 0.24 | 11 | 0.00044047 |  | 2.1484 | 0.331646254349767-0.114672968202751i | 0.53797 | TRUE | 2.1484
| 0.26 | 12 | -0.0010085 |  | 2.3438 | -0.246005780246763+0.188028128757247i | 0.47469 | TRUE | 2.3438
| 0.28 | 13 | 0.00086031 |  | 2.5391 | 0.191057871957205-0.241941953898917i | 0.47262 | TRUE | 2.5391
| 0.3 | 14 | 0.00049689 |  | 2.7344 | -0.103716355130832+0.2820540888195i | 0.46071 | TRUE | 2.7344
| 0.32 | 15 | -0.0005061 |  | 2.9297 | 2.5334588576684E-002-0.292203244137455i | 0.44964 | TRUE | 2.9297
| 0.34 | 16 | 0.00046484 |  | 3.125 | 6.74951549936766E-002+0.277765151776836i | 0.43822 | TRUE | 3.125
| 0.36 | 17 | -0.00035678 |  | 3.3203 | -0.150538031066545-0.260976277726258i | 0.46188 | TRUE | 3.3203
| 0.38 | 18 | 9.4809e-07 |  | 3.5156 | 0.241715818071617+0.170531409634939i | 0.4535 | TRUE | 3.5156
| 0.4 | 19 | -0.00040691 |  | 3.7109 | -0.237851891554428-8.16021987163241E-002i | 0.3855 | TRUE | 3.7109
| 0.42 | 20 | 7.1603e-05 |  | 3.9062 | 0.261910863599709+1.73083542965981E-002i | 0.4024 | TRUE | 3.9062
| 0.44 | 21 | 0.00071227 |  | 4.1016 | -0.22163855808993+7.29289896982082E-002i | 0.35771 | TRUE | 4.1016
| 0.46 | 22 | 0.00049424 |  | 4.2969 | 0.189935140226876-0.122453976894655i | 0.34645 | TRUE | 4.2969
| 0.48 | 23 | -0.0011427 |  | 4.4922 | -0.133077751670883+0.170833566528893i | 0.33198 | TRUE | 4.4922
| 0.5 | 24 | 0.00071924 |  | 4.6875 | 8.6955863055297E-002-0.19631894596465i | 0.32917 | TRUE | 4.6875
| 0.52 | 25 | -0.00014438 |  | 4.8828 | -6.09666699275283E-003+0.196122847048429i | 0.30081 | TRUE | 4.8828
| 0.54 | 26 | -0.00027556 |  | 5.0781 | -3.52336414338883E-002-0.186416444036064i | 0.29085 | TRUE | 5.0781
| 0.56 | 27 | 0.00029271 |  | 5.2734 | 9.13697555643868E-002+0.152531314398193i | 0.27258 | TRUE | 5.2734
| 0.58 | 28 | 0.00027109 |  | 5.4688 | -9.04677210320092E-002-0.118750128370904i | 0.22886 | TRUE | 5.4688
| 0.6 | 29 | -0.00070469 |  | 5.6641 | 0.175543233808119+0.129259108842603i | 0.3342 | TRUE | 5.6641
| 0.62 | 30 | -0.00012813 |  | 5.8594 | -0.205806031037833-1.7589809849604E-002i | 0.31666 | TRUE | 5.8594
| 0.64 | 31 | 0.00064464 |  | 6.0547 | 0.179666632698211-5.10062367696549E-002i | 0.28632 | TRUE | 6.0547
| 0.66 | 32 | -0.00078271 |  | 6.25 | -0.13569207659537+0.115919029531062i | 0.2736 | TRUE | 6.25
| 0.68 | 33 | 0.00019075 |  | 6.4453 | 0.1019820684054-0.111072854231237i | 0.23117 | TRUE | 6.4453
| 0.7 | 34 | 0.00047149 |  | 6.6406 | -7.28877179390679E-002+0.149209165321778i | 0.25458 | TRUE | 6.6406
| 0.72 | 35 | 0.00086263 |  | 6.8359 | 2.01835783940895E-002-0.14973932262311i | 0.23164 | TRUE | 6.8359
| 0.74 | 36 | -0.00063342 |  | 7.0312 | 4.02201274257722E-002+0.160445341821829i | 0.25358 | TRUE | 7.0312
| 0.76 | 37 | -0.00015577 |  | 7.2266 | -6.85428898786092E-002-0.120186536942258i | 0.21211 | TRUE | 7.2266
| 0.78 | 38 | 0.00038392 |  | 7.4219 | 0.107523493753152+0.123436120977857i | 0.25096 | TRUE | 7.4219
| 0.8 | 39 | 0.00010201 |  | 7.6172 | -0.152081664975886-5.97638588113851E-002i | 0.25051 | TRUE | 7.6172
| 0.82 | 40 | -0.00053624 |  | 7.8125 | 0.131691887999408-2.03284075610054E-003i | 0.20192 | TRUE | 7.8125
| 0.84 | 41 | -0.0005883 |  | 8.0078 | -0.13085321451555+2.15947594901312E-002i | 0.20332 | TRUE | 8.0078
| 0.86 | 42 | 0.00088275 |  | 8.2031 | 0.138372556657144-7.37503306164103E-002i | 0.24038 | TRUE | 8.2031
| 0.88 | 43 | 0.00050872 |  | 8.3984 | -6.79722870933307E-002+0.126538384292797i | 0.22021 | TRUE | 8.3984
| 0.9 | 44 | -0.00049181 |  | 8.5938 | 5.05715877455992E-002-9.95711525925625E-002i | 0.17121 | TRUE | 8.5938
| 0.92 | 45 | -8.067e-05 |  | 8.7891 | -1.40272617101881E-002+0.136067922327444i | 0.20971 | TRUE | 8.7891
| 0.94 | 46 | 5.3295e-05 |  | 8.9844 | -4.15284504054995E-002-7.69343540552364E-002i | 0.13403 | TRUE | 8.9844
| 0.96 | 47 | 0.00017612 |  | 9.1797 | 4.21766321356077E-002+0.101020869771092i | 0.16783 | TRUE | 9.1797
| 0.98 | 48 | -0.00015682 |  | 9.375 | -7.82709784747372E-002-8.23929156648509E-002i | 0.17422 | TRUE | 9.375
| 1 | 49 | 0.00045568 |  | 9.5703 | 9.52028065227263E-002+5.21566820471817E-002i | 0.16642 | TRUE | 9.5703
| 1.02 | 50 | -0.00010834 |  | 9.7656 | -9.51097154048865E-002-9.98361436644265E-003i | 0.14661 | TRUE | 9.7656
| 1.04 | 51 | -0.00014422 |  | 9.9609 | 0.103133930545341+3.11187981677839E-003i | 0.15818 | TRUE | 9.9609
| 1.06 | 52 | -0.00056851 |  | 10.1562 | -9.68566437367716E-002+4.90034146104565E-002i | 0.16641 | TRUE | 10.1562
| 1.08 | 53 | 0.0022465 |  | 10.3516 | 7.01264115061105E-002-4.62650373999717E-002i | 0.1288 | TRUE | 10.3516
| 1.1 | 54 | -0.00043885 |  | 10.5469 | -5.86620170140824E-002+8.71311677822379E-002i | 0.16103 | TRUE | 10.5469
| 1.12 | 55 | -0.00049084 |  | 10.7422 | 1.65569933520599E-002-8.90865613265036E-002i | 0.13891 | TRUE | 10.7422
| 1.14 | 56 | 0.0007118 |  | 10.9375 | -8.63200798534962E-003+9.00734688983441E-002i | 0.13872 | TRUE | 10.9375
| 1.16 | 57 | -0.00029933 |  | 11.1328 | -4.38643952916442E-002-8.41037051136162E-002i | 0.14542 | TRUE | 11.1328
| 1.18 | 58 | -0.00087608 |  | 11.3281 | 6.37303940023734E-002+8.66482580400961E-002i | 0.1649 | TRUE | 11.3281
| 1.2 | 59 | 0.00012943 |  | 11.5234 | -9.51429503526897E-002-4.81306178114482E-002i | 0.16346 | TRUE | 11.5234
| 1.22 | 60 | 0.00033429 |  | 11.7188 | 9.51298571046461E-002+6.73951297079519E-003i | 0.14621 | TRUE | 11.7188
| 1.24 | 61 | 0.00050198 |  | 11.9141 | -8.7070539463692E-002+2.78237248733497E-003i | 0.13355 | TRUE | 11.9141
| 1.26 | 62 | 0.00056541 |  | 12.1094 | 6.15577909853808E-002-3.02605005804621E-002i | 0.10516 | TRUE | 12.1094
| 1.28 | 63 | -0.0007567 |  | 12.3047 | -4.63547098451829E-002+3.15423247515648E-002i | 0.085956 | TRUE | 12.3047
| 1.3 | 64 | 0.0009418 |  | 12.5 | 3.22572921649999E-002-6.77598564249996E-002i | 0.11505 | TRUE | 12.5
| 1.32 | 65 | -0.00017716 |  | 12.6953 | -6.05554715634981E-003+8.00083639065715E-002i | 0.12301 | TRUE | 12.6953
| 1.34 | 66 | 0.0011509 |  | 12.8906 | -2.14560639400528E-002-6.780624088015E-002i | 0.10903 | TRUE | 12.8906
| 1.36 | 67 | -0.00098081 |  | 13.0859 | 4.99280820278143E-002+6.1392991315015E-002i | 0.12131 | TRUE | 13.0859
| 1.38 | 68 | -0.000496 |  | 13.2812 | -4.63585947505337E-002-3.88079024983861E-002i | 0.092686 | TRUE | 13.2812
| 1.4 | 69 | 0.00058831 |  | 13.4766 | 3.70210135448826E-002+2.85621504948567E-002i | 0.071683 | TRUE | 13.4766
| 1.42 | 70 | -6.889e-06 |  | 13.6719 | -8.31189891000107E-002-2.23889819603655E-002i | 0.13197 | TRUE | 13.6719
| 1.44 | 71 | 0.00084159 |  | 13.8672 | 7.41155085966581E-002-2.18347992014445E-002i | 0.11845 | TRUE | 13.8672
| 1.46 | 72 | 0.00010577 |  | 14.0625 | -2.50898868536337E-002+5.33826437890365E-002i | 0.090427 | TRUE | 14.0625
| 1.48 | 73 | -0.00053965 |  | 14.2578 | 4.48041775045546E-002-1.98380154247536E-002i | 0.075119 | TRUE | 14.2578
| 1.5 | 74 | 0.00045933 |  | 14.4531 | -2.55059364838721E-002+4.22051505269931E-002i | 0.075601 | TRUE | 14.4531
| 1.52 | 75 | -0.0011225 |  | 14.6484 | 4.24852210983642E-003-3.89140751837866E-002i | 0.060012 | TRUE | 14.6484
| 1.54 | 76 | 0.001237 |  | 14.8438 | -1.34876996102625E-002+3.04245177815613E-002i | 0.05102 | TRUE | 14.8438
| 1.56 | 77 | -0.00080321 |  | 15.0391 | -1.38881067857513E-002-4.74312948527044E-002i | 0.075768 | TRUE | 15.0391
| 1.58 | 78 | 0.00073699 |  | 15.2344 | 3.04718742541213E-002+1.98240378416233E-002i | 0.055731 | TRUE | 15.2344
| 1.6 | 79 | -0.00016548 |  | 15.4297 | -3.86830633880153E-002-3.90479269230339E-002i | 0.084264 | TRUE | 15.4297
| 1.62 | 80 | -0.00032477 |  | 15.625 | 4.27535881527952E-002+1.69899830867296E-002i | 0.070529 | TRUE | 15.625
| 1.64 | 81 | 0.0010507 |  | 15.8203 | -5.20432392637507E-002-2.55567578656877E-003i | 0.079881 | TRUE | 15.8203
| 1.66 | 82 | -0.00036492 |  | 16.0156 | 5.66066230540693E-002-2.36652362075726E-002i | 0.09406 | TRUE | 16.0156
| 1.68 | 83 | 0.00015701 |  | 16.2109 | -9.11861769430376E-003+2.36603542841727E-002i | 0.038873 | TRUE | 16.2109
| 1.7 | 84 | -0.00010423 |  | 16.4062 | 2.76100744911676E-002-3.41393854241812E-002i | 0.067312 | TRUE | 16.4062
| 1.72 | 85 | 0.00011027 |  | 16.6016 | 1.67306133123092E-003+6.00445466377385E-002i | 0.092087 | TRUE | 16.6016
| 1.74 | 86 | 0.001034 |  | 16.7969 | -1.55983781425025E-002-3.90618643479299E-002i | 0.064482 | TRUE | 16.7969
| 1.76 | 87 | 0.00053486 |  | 16.9922 | 2.66167230699446E-002+3.13661862187519E-002i | 0.063066 | TRUE | 16.9922
| 1.78 | 88 | 0.0002695 |  | 17.1875 | -1.594857730446E-002-1.42364253987861E-002i | 0.032774 | TRUE | 17.1875
| 1.8 | 89 | 9.1227e-05 |  | 17.3828 | 1.29484847078959E-002+4.94021861893539E-002i | 0.078295 | TRUE | 17.3828
| 1.82 | 90 | -0.00020726 |  | 17.5781 | -5.79195251240572E-002-1.79157383819419E-002i | 0.092945 | TRUE | 17.5781
| 1.84 | 91 | -0.00020952 |  | 17.7734 | 6.71458020961027E-002+2.62899961880245E-003i | 0.10302 | TRUE | 17.7734
| 1.86 | 92 | 0.0001792 |  | 17.9688 | -1.21158803573514E-002+4.85305748521686E-002i | 0.076684 | TRUE | 17.9688
| 1.88 | 93 | 0.00055363 |  | 18.1641 | 9.16936470187561E-003+5.08491013652895E-003i | 0.016074 | TRUE | 18.1641
| 1.9 | 94 | 0.00052841 |  | 18.3594 | -3.08035399810569E-002+3.99751273141408E-002i | 0.077368 | TRUE | 18.3594
| 1.92 | 95 | -3.2729e-05 |  | 18.5547 | 1.10121986037159E-003-3.05339311466259E-002i | 0.046841 | TRUE | 18.5547
| 1.94 | 96 | 8.2515e-05 |  | 18.75 | 1.96012828753696E-002+4.74242813210622E-002i | 0.078669 | TRUE | 18.75
| 1.96 | 97 | -0.00030422 |  | 18.9453 | -3.28348255122368E-002-1.83868803770725E-002i | 0.057693 | TRUE | 18.9453
| 1.98 | 98 | 0.001039 |  | 19.1406 | 2.14407330821194E-002+1.48677089051729E-003i | 0.032949 | TRUE | 19.1406
| 2 | 99 | 0.0011359 |  | 19.3359 | -3.03456112283158E-002+3.12196611101287E-003i | 0.046767 | TRUE | 19.3359
| 2.02 | 100 | -9.5898e-05 |  | 19.5312 | -3.71771139591313E-003+6.50197328814041E-003i | 0.011482 | TRUE | 19.5312
| 2.04 | 101 | 0.00032387 |  | 19.7266 | -2.66991184344792E-003-1.1979942102943E-002i | 0.018817 | TRUE | 19.7266
| 2.06 | 102 | 0.0013493 |  | 19.9219 | 6.9922435176673E-003-6.65456702099547E-003i | 0.014798 | TRUE | 19.9219
| 2.08 | 103 | 0.00041517 |  | 20.1172 | -4.11661259100704E-002+1.38053872428547E-002i | 0.066564 | TRUE | 20.1172
| 2.1 | 104 | -0.00036413 |  | 20.3125 | 1.07127699626409E-002-3.39321866659415E-002i | 0.054551 | TRUE | 20.3125
| 2.12 | 105 | 0.00064193 |  | 20.5078 | 1.48836722121721E-002+2.2987660894377E-002i | 0.041983 | TRUE | 20.5078
| 2.14 | 106 | 0.0013984 |  | 20.7031 | -3.91658721163961E-002-1.5753354711562E-002i | 0.064719 | TRUE | 20.7031
| 2.16 | 107 | 0.00012676 |  | 20.8984 | 1.4341604269413E-002+8.21817563357587E-003i | 0.02534 | TRUE | 20.8984
| 2.18 | 108 | 0.00013569 |  | 21.0938 | -2.27651843979862E-003-5.11449400332691E-003i | 0.0085825 | TRUE | 21.0938
| 2.2 | 109 | 0.00078329 |  | 21.2891 | -8.20018102610333E-003+7.07685333498818E-003i | 0.016606 | TRUE | 21.2891
| 2.22 | 110 | 0.0015033 |  | 21.4844 | -2.29169660664395E-002-1.73974913444555E-002i | 0.04411 | TRUE | 21.4844
| 2.24 | 111 | 0.00079303 |  | 21.6797 | 1.85473481100979E-002-3.03710493685436E-003i | 0.028813 | TRUE | 21.6797
| 2.26 | 112 | 0.0013856 |  | 21.875 | -6.3468839917347E-003-1.18929974315829E-002i | 0.020667 | TRUE | 21.875
| 2.28 | 113 | -0.00085115 |  | 22.0703 | 6.57147660720657E-003-1.47145666264112E-003i | 0.010324 | TRUE | 22.0703
| 2.3 | 114 | 0.0012398 |  | 22.2656 | -1.21228658681938E-002+2.53566702248209E-002i | 0.043087 | TRUE | 22.2656
| 2.32 | 115 | 0.0018217 |  | 22.4609 | 2.24953126335076E-003-2.01687767501535E-002i | 0.031112 | TRUE | 22.4609
| 2.34 | 116 | -7.5735e-06 |  | 22.6562 | 1.80219070496041E-002+9.03759821831979E-003i | 0.030908 | TRUE | 22.6562
| 2.36 | 117 | 0.0011838 |  | 22.8516 | 8.80437017428287E-003-1.54351374718885E-002i | 0.027242 | TRUE | 22.8516
| 2.38 | 118 | 0.0012015 |  | 23.0469 | 1.87209278357088E-002+3.06925748263665E-002i | 0.055116 | TRUE | 23.0469
| 2.4 | 119 | 0.002336 |  | 23.2422 | -4.86432254169837E-002+7.02268085758308E-003i | 0.075346 | TRUE | 23.2422
| 2.42 | 120 | 0.0017542 |  | 23.4375 | 1.2913561513187E-002-1.45117741955645E-002i | 0.02978 | TRUE | 23.4375
| 2.44 | 121 | 0.0010822 |  | 23.6328 | -1.00668300134416E-002+1.83528944284308E-002i | 0.032091 | TRUE | 23.6328
| 2.46 | 122 | 0.0026145 |  | 23.8281 | -2.190166640727E-002-1.74227251939578E-002i | 0.042905 | TRUE | 23.8281
| 2.48 | 123 | 0.0029112 |  | 24.0234 | -1.20731422643458E-002+8.39272691132392E-003i | 0.022542 | TRUE | 24.0234
| 2.5 | 124 | 0.0031632 |  | 24.2188 | 8.39199694554831E-003-2.18503355969624E-003i | 0.013294 | TRUE | 24.2188
| 2.52 | 125 | 0.0028333 |  | 24.4141 | 2.96556320189795E-004+3.23100760222383E-002i | 0.049535 | TRUE | 24.4141
| 2.54 | 126 | 0.0053121 |  | 24.6094 | -3.72048957646896E-002-1.56199887674683E-002i | 0.06186 | TRUE | 24.6094
| 2.56 | 127 | 0.0051282 |  | 24.8047 | 1.68031139705577E-003-5.7404360147012E-003i | 0.0091697 | TRUE | 24.8047
| 2.58 | 128 | 0.0047971 |  | 25 | -2.64172717999978E-003 | 0.0040499 | TRUE | 25
| 2.6 | 129 | 0.006384 |  | 25.1953 | 1.68031139705554E-003+5.74043601470081E-003i | 0 | FALSE | 25.1953
| 2.62 | 130 | 0.006327 |  | 25.3906 | -3.72048957646892E-002+1.56199887674688E-002i | 0 | FALSE | 25.3906
| 2.64 | 131 | 0.0063802 |  | 25.5859 | 2.96556320189267E-004-3.23100760222384E-002i | 0 | FALSE | 25.5859
| 2.66 | 132 | 0.0069704 |  | 25.7812 | 8.39199694554862E-003+2.18503355969632E-003i | 0 | FALSE | 25.7812
| 2.68 | 133 | 0.0078436 |  | 25.9766 | -1.20731422643462E-002-8.39272691132384E-003i | 0 | FALSE | 25.9766
| 2.7 | 134 | 0.0098536 |  | 26.1719 | -2.19016664072697E-002+1.74227251939578E-002i | 0 | FALSE | 26.1719
| 2.72 | 135 | 0.013412 |  | 26.3672 | -1.00668300134419E-002-1.83528944284305E-002i | 0 | FALSE | 26.3672
| 2.74 | 136 | 0.015918 |  | 26.5625 | 1.29135615131871E-002+1.45117741955643E-002i | 0 | FALSE | 26.5625
| 2.76 | 137 | 0.020558 |  | 26.7578 | -4.86432254169838E-002-7.02268085758259E-003i | 0 | FALSE | 26.7578
| 2.78 | 138 | 0.028303 |  | 26.9531 | 1.87209278357086E-002-3.06925748263668E-002i | 0 | FALSE | 26.9531
| 2.8 | 139 | 0.048061 |  | 27.1484 | 8.80437017428312E-003+1.54351374718886E-002i | 0 | FALSE | 27.1484
| 2.82 | 140 | 0.099524 |  | 27.3438 | 1.8021907049604E-002-9.03759821831998E-003i | 0 | FALSE | 27.3438
| 2.84 | 141 | 0.12194 |  | 27.5391 | 2.24953126335101E-003+2.01687767501536E-002i | 0 | FALSE | 27.5391
| 2.86 | 142 | 0.060366 |  | 27.7344 | -1.21228658681942E-002-2.53566702248209E-002i | 0 | FALSE | 27.7344
| 2.88 | 143 | 0.03279 |  | 27.9297 | 6.57147660720679E-003+1.47145666264104E-003i | 0 | FALSE | 27.9297
| 2.9 | 144 | 0.022808 |  | 28.125 | -6.34688399173482E-003+1.1892997431583E-002i | 0 | FALSE | 28.125
| 2.92 | 145 | 0.017193 |  | 28.3203 | 1.8547348110098E-002+3.03710493685411E-003i | 0 | FALSE | 28.3203
| 2.94 | 146 | 0.013076 |  | 28.5156 | -2.29169660664396E-002+1.73974913444558E-002i | 0 | FALSE | 28.5156
| 2.96 | 147 | 0.011186 |  | 28.7109 | -8.20018102610333E-003-7.07685333498823E-003i | 0 | FALSE | 28.7109
| 2.98 | 148 | 0.0075542 |  | 28.9062 | -2.27651843979848E-003+5.11449400332712E-003i | 0 | FALSE | 28.9062
| 3 | 149 | 0.010433 |  | 29.1016 | 1.43416042694129E-002-8.21817563357597E-003i | 0 | FALSE | 29.1016
| 3.02 | 150 | 0.0051782 |  | 29.2969 | -3.9165872116396E-002+1.57533547115623E-002i | 0 | FALSE | 29.2969
| 3.04 | 151 | 0.0011919 |  | 29.4922 | 1.48836722121719E-002-2.29876608943771E-002i | 0 | FALSE | 29.4922
| 3.06 | 152 | 0.0040349 |  | 29.6875 | 1.07127699626411E-002+3.39321866659414E-002i | 0 | FALSE | 29.6875
| 3.08 | 153 | 0.0045647 |  | 29.8828 | -4.11661259100706E-002-1.38053872428545E-002i | 0 | FALSE | 29.8828
| 3.1 | 154 | 0.0059729 |  | 30.0781 | 6.99224351766738E-003+6.65456702099539E-003i | 0 | FALSE | 30.0781
| 3.12 | 155 | 0.0033597 |  | 30.2734 | -2.66991184344789E-003+1.19799421029431E-002i | 0 | FALSE | 30.2734
| 3.14 | 156 | 0.0039217 |  | 30.4688 | -3.71771139591313E-003-6.50197328814035E-003i | 0 | FALSE | 30.4688
| 3.16 | 157 | 0.0021422 |  | 30.6641 | -3.03456112283159E-002-3.12196611101256E-003i | 0 | FALSE | 30.6641
| 3.18 | 158 | 0.003002 |  | 30.8594 | 2.14407330821193E-002-1.48677089051746E-003i | 0 | FALSE | 30.8594
| 3.2 | 159 | 0.001599 |  | 31.0547 | -3.28348255122367E-002+1.83868803770727E-002i | 0 | FALSE | 31.0547
| 3.22 | 160 | 0.0021704 |  | 31.25 | 1.96012828753695E-002-4.74242813210623E-002i | 0 | FALSE | 31.25
| 3.24 | 161 | -0.0022857 |  | 31.4453 | 1.10121986037161E-003+3.05339311466259E-002i | 0 | FALSE | 31.4453
| 3.26 | 162 | -0.0018575 |  | 31.6406 | -3.08035399810571E-002-3.99751273141408E-002i | 0 | FALSE | 31.6406
| 3.28 | 163 | 0.0022665 |  | 31.8359 | 9.16936470187557E-003-5.08491013652898E-003i | 0 | FALSE | 31.8359
| 3.3 | 164 | -0.0017756 |  | 32.0312 | -1.21158803573517E-002-4.85305748521685E-002i | 0 | FALSE | 32.0312
| 3.32 | 165 | 0.0034271 |  | 32.2266 | 6.71458020961027E-002-2.62899961880276E-003i | 0 | FALSE | 32.2266
| 3.34 | 166 | 0.0036274 |  | 32.4219 | -5.79195251240571E-002+1.79157383819421E-002i | 0 | FALSE | 32.4219
| 3.36 | 167 | 0.003039 |  | 32.6172 | 1.29484847078957E-002-4.9402186189354E-002i | 0 | FALSE | 32.6172
| 3.38 | 168 | 0.0023565 |  | 32.8125 | -1.594857730446E-002+1.42364253987862E-002i | 0 | FALSE | 32.8125
| 3.4 | 169 | 0.00035578 |  | 33.0078 | 2.66167230699444E-002-3.13661862187521E-002i | 0 | FALSE | 33.0078
| 3.42 | 170 | 0.0012374 |  | 33.2031 | -1.55983781425024E-002+3.90618643479299E-002i | 0 | FALSE | 33.2031
| 3.44 | 171 | 0.0013582 |  | 33.3984 | 1.67306133123069E-003-6.00445466377386E-002i | 0 | FALSE | 33.3984
| 3.46 | 172 | 0.0014448 |  | 33.5938 | 2.76100744911677E-002+3.41393854241812E-002i | 0 | FALSE | 33.5938
| 3.48 | 173 | 0.0036312 |  | 33.7891 | -9.11861769430388E-003-2.36603542841728E-002i | 0 | FALSE | 33.7891
| 3.5 | 174 | -0.00029147 |  | 33.9844 | 5.66066230540693E-002+2.36652362075724E-002i | 0 | FALSE | 33.9844
| 3.52 | 175 | 0.0041889 |  | 34.1797 | -5.20432392637507E-002+2.55567578656884E-003i | 0 | FALSE | 34.1797
| 3.54 | 176 | 0.00051689 |  | 34.375 | 4.27535881527953E-002-1.69899830867297E-002i | 0 | FALSE | 34.375
| 3.56 | 177 | 0.00069165 |  | 34.5703 | -3.86830633880152E-002+3.9047926923034E-002i | 0 | FALSE | 34.5703
| 3.58 | 178 | -0.00032115 |  | 34.7656 | 3.04718742541214E-002-1.98240378416234E-002i | 0 | FALSE | 34.7656
| 3.6 | 179 | -0.0045134 |  | 34.9609 | -1.38881067857512E-002+4.74312948527044E-002i | 0 | FALSE | 34.9609
| 3.62 | 180 | 0.0025142 |  | 35.1562 | -1.34876996102624E-002-3.04245177815613E-002i | 0 | FALSE | 35.1562
| 3.64 | 181 | 0.00019789 |  | 35.3516 | 4.24852210983649E-003+3.89140751837867E-002i | 0 | FALSE | 35.3516
| 3.66 | 182 | 0.00066274 |  | 35.5469 | -2.55059364838723E-002-4.22051505269932E-002i | 0 | FALSE | 35.5469
| 3.68 | 183 | 0.0033253 |  | 35.7422 | 4.48041775045546E-002+1.98380154247536E-002i | 0 | FALSE | 35.7422
| 3.7 | 184 | 0.0033608 |  | 35.9375 | -2.50898868536339E-002-5.33826437890366E-002i | 0 | FALSE | 35.9375
| 3.72 | 185 | 5.4657e-05 |  | 36.1328 | 7.41155085966583E-002+2.18347992014443E-002i | 0 | FALSE | 36.1328
| 3.74 | 186 | 0.00096285 |  | 36.3281 | -8.31189891000108E-002+2.23889819603657E-002i | 0 | FALSE | 36.3281
| 3.76 | 187 | -0.0026011 |  | 36.5234 | 3.70210135448827E-002-2.85621504948567E-002i | 0 | FALSE | 36.5234
| 3.78 | 188 | 0.0023586 |  | 36.7188 | -4.63585947505338E-002+3.88079024983863E-002i | 0 | FALSE | 36.7188
| 3.8 | 189 | -0.00038922 |  | 36.9141 | 4.99280820278143E-002-6.13929913150154E-002i | 0 | FALSE | 36.9141
| 3.82 | 190 | -0.00067824 |  | 37.1094 | -2.14560639400529E-002+6.78062408801502E-002i | 0 | FALSE | 37.1094
| 3.84 | 191 | -0.0003534 |  | 37.3047 | -6.0555471563499E-003-8.00083639065718E-002i | 0 | FALSE | 37.3047
| 3.86 | 192 | 0.0043834 |  | 37.5 | 3.2257292165E-002+6.77598564249996E-002i | 0 | FALSE | 37.5
| 3.88 | 193 | -0.0023066 |  | 37.6953 | -4.63547098451829E-002-3.15423247515649E-002i | 0 | FALSE | 37.6953
| 3.9 | 194 | 0.0013642 |  | 37.8906 | 6.1557790985381E-002+3.02605005804622E-002i | 0 | FALSE | 37.8906
| 3.92 | 195 | -0.00016326 |  | 38.0859 | -8.70705394636921E-002-2.7823724873348E-003i | 0 | FALSE | 38.0859
| 3.94 | 196 | -0.0034759 |  | 38.2812 | 9.51298571046462E-002-6.73951297079534E-003i | 0 | FALSE | 38.2812
| 3.96 | 197 | -0.00093371 |  | 38.4766 | -9.51429503526897E-002+4.81306178114485E-002i | 0 | FALSE | 38.4766
| 3.98 | 198 | -0.00074196 |  | 38.6719 | 6.37303940023731E-002-8.66482580400963E-002i | 0 | FALSE | 38.6719
| 4 | 199 | 0.00019634 |  | 38.8672 | -4.3864395291644E-002+8.41037051136164E-002i | 0 | FALSE | 38.8672
| 4.02 | 200 | 0.0021781 |  | 39.0625 | -8.63200798534982E-003-9.00734688983442E-002i | 0 | FALSE | 39.0625
| 4.04 | 201 | 0.0005306 |  | 39.2578 | 1.65569933520602E-002+8.90865613265037E-002i | 0 | FALSE | 39.2578
| 4.06 | 202 | -0.00027096 |  | 39.4531 | -5.86620170140827E-002-8.71311677822378E-002i | 0 | FALSE | 39.4531
| 4.08 | 203 | 0.00085068 |  | 39.6484 | 7.01264115061107E-002+4.62650373999716E-002i | 0 | FALSE | 39.6484
| 4.1 | 204 | -0.0023017 |  | 39.8438 | -9.68566437367719E-002-4.90034146104562E-002i | 0 | FALSE | 39.8438
| 4.12 | 205 | -0.0026624 |  | 40.0391 | 0.103133930545341-3.1118798167787E-003i | 0 | FALSE | 40.0391
| 4.14 | 206 | 0.0016838 |  | 40.2344 | -9.51097154048866E-002+9.98361436644301E-003i | 0 | FALSE | 40.2344
| 4.16 | 207 | -0.00097513 |  | 40.4297 | 9.52028065227263E-002-5.21566820471822E-002i | 0 | FALSE | 40.4297
| 4.18 | 208 | 0.0011405 |  | 40.625 | -7.82709784747371E-002+8.23929156648512E-002i | 0 | FALSE | 40.625
| 4.2 | 209 | 6.0797e-06 |  | 40.8203 | 4.21766321356074E-002-0.101020869771092i | 0 | FALSE | 40.8203
| 4.22 | 210 | 0.0019961 |  | 41.0156 | -4.15284504054993E-002+7.69343540552368E-002i | 0 | FALSE | 41.0156
| 4.24 | 211 | -0.0010846 |  | 41.2109 | -1.40272617101887E-002-0.136067922327444i | 0 | FALSE | 41.2109
| 4.26 | 212 | 4.8816e-05 |  | 41.4062 | 5.05715877455995E-002+9.95711525925625E-002i | 0 | FALSE | 41.4062
| 4.28 | 213 | 0.0014499 |  | 41.6016 | -6.79722870933313E-002-0.126538384292797i | 0 | FALSE | 41.6016
| 4.3 | 214 | 0.002875 |  | 41.7969 | 0.138372556657144+7.37503306164097E-002i | 0 | FALSE | 41.7969
| 4.32 | 215 | 0.00029469 |  | 41.9922 | -0.13085321451555-2.15947594901308E-002i | 0 | FALSE | 41.9922
| 4.34 | 216 | 0.00046098 |  | 42.1875 | 0.131691887999409+2.03284075609992E-003i | 0 | FALSE | 42.1875
| 4.36 | 217 | 0.0012168 |  | 42.3828 | -0.152081664975885+5.97638588113858E-002i | 0 | FALSE | 42.3828
| 4.38 | 218 | -0.0017959 |  | 42.5781 | 0.107523493753151-0.123436120977858i | 0 | FALSE | 42.5781
| 4.4 | 219 | 0.0011011 |  | 42.7734 | -6.85428898786088E-002+0.120186536942259i | 0 | FALSE | 42.7734
| 4.42 | 220 | 0.0025125 |  | 42.9688 | 4.02201274257714E-002-0.160445341821829i | 0 | FALSE | 42.9688
| 4.44 | 221 | 0.00049525 |  | 43.1641 | 2.01835783940903E-002+0.14973932262311i | 0 | FALSE | 43.1641
| 4.46 | 222 | -0.002773 |  | 43.3594 | -7.28877179390687E-002-0.149209165321778i | 0 | FALSE | 43.3594
| 4.48 | 223 | -0.0031417 |  | 43.5547 | 0.1019820684054+0.111072854231237i | 0 | FALSE | 43.5547
| 4.5 | 224 | -0.0018121 |  | 43.75 | -0.13569207659537-0.115919029531062i | 0 | FALSE | 43.75
| 4.52 | 225 | 0.0011027 |  | 43.9453 | 0.179666632698211+5.1006236769654E-002i | 0 | FALSE | 43.9453
| 4.54 | 226 | 8.7023e-05 |  | 44.1406 | -0.205806031037834+1.7589809849605E-002i | 0 | FALSE | 44.1406
| 4.56 | 227 | 0.0016575 |  | 44.3359 | 0.175543233808119-0.129259108842605i | 0 | FALSE | 44.3359
| 4.58 | 228 | 0.00066954 |  | 44.5312 | -9.04677210320088E-002+0.118750128370904i | 0 | FALSE | 44.5312
| 4.6 | 229 | -0.0010286 |  | 44.7266 | 9.13697555643862E-002-0.152531314398193i | 0 | FALSE | 44.7266
| 4.62 | 230 | -0.0010079 |  | 44.9219 | -3.52336414338876E-002+0.186416444036064i | 0 | FALSE | 44.9219
| 4.64 | 231 | 0.0018722 |  | 45.1172 | -6.09666699275369E-003-0.196122847048429i | 0 | FALSE | 45.1172
| 4.66 | 232 | 0.00031453 |  | 45.3125 | 8.6955863055298E-002+0.196318945964649i | 0 | FALSE | 45.3125
| 4.68 | 233 | -0.001469 |  | 45.5078 | -0.133077751670884-0.170833566528893i | 0 | FALSE | 45.5078
| 4.7 | 234 | -9.8109e-05 |  | 45.7031 | 0.189935140226876+0.122453976894654i | 0 | FALSE | 45.7031
| 4.72 | 235 | -0.0011725 |  | 45.8984 | -0.22163855808993-7.29289896982073E-002i | 0 | FALSE | 45.8984
| 4.74 | 236 | 0.00054139 |  | 46.0938 | 0.261910863599709-1.73083542965993E-002i | 0 | FALSE | 46.0938
| 4.76 | 237 | 0.00064778 |  | 46.2891 | -0.237851891554428+8.16021987163252E-002i | 0 | FALSE | 46.2891
| 4.78 | 238 | 0.0015155 |  | 46.4844 | 0.241715818071617-0.170531409634941i | 0 | FALSE | 46.4844
| 4.8 | 239 | -8.4579e-05 |  | 46.6797 | -0.150538031066544+0.260976277726259i | 0 | FALSE | 46.6797
| 4.82 | 240 | 0.0001933 |  | 46.875 | 6.74951549936758E-002-0.277765151776837i | 0 | FALSE | 46.875
| 4.84 | 241 | 0.00057071 |  | 47.0703 | 2.53345885766854E-002+0.292203244137455i | 0 | FALSE | 47.0703
| 4.86 | 242 | -0.0012555 |  | 47.2656 | -0.103716355130834-0.2820540888195i | 0 | FALSE | 47.2656
| 4.88 | 243 | -0.00071436 |  | 47.4609 | 0.191057871957206+0.241941953898916i | 0 | FALSE | 47.4609
| 4.9 | 244 | 0.0022515 |  | 47.6562 | -0.246005780246764-0.188028128757246i | 0 | FALSE | 47.6562
| 4.92 | 245 | 0.0015814 |  | 47.8516 | 0.331646254349768+0.11467296820275i | 0 | FALSE | 47.8516
| 4.94 | 246 | -0.0016907 |  | 48.0469 | -0.359205252998011+1.8854911545878E-002i | 0 | FALSE | 48.0469
| 4.96 | 247 | 0.0018917 |  | 48.2422 | 0.347682175679155-0.120361615607488i | 0 | FALSE | 48.2422
| 4.98 | 248 | -0.0021461 |  | 48.4375 | -0.322576769907088+0.240964966227652i | 0 | FALSE | 48.4375
| 5 | 249 | 0.00042349 |  | 48.6328 | 0.239691480199042-0.366617529333173i | 0 | FALSE | 48.6328
| 5.02 | 250 | 6.4552e-05 |  | 48.8281 | -0.105270143894247+0.44236257106016i | 0 | FALSE | 48.8281
| 5.04 | 251 | -0.0034784 |  | 49.0234 | -2.44411015305457E-002-0.471563541940732i | 0 | FALSE | 49.0234
| 5.06 | 252 | -0.0015648 |  | 49.2188 | 0.176858839367356+0.475634274077495i | 0 | FALSE | 49.2188
| 5.08 | 253 | 0.0016376 |  | 49.4141 | -0.328773322296171-0.439889608747543i | 0 | FALSE | 49.4141
| 5.1 | 254 | 0.0010824 |  | 49.6094 | 0.491167230092665+0.35123215589877i | 0 | FALSE | 49.6094
| 5.12 | 255 | 0.0042539 |  | 49.8047 | -0.6204026677677-0.201455622694085i | 0 | FALSE | 49.8047
Note on I4: Luca Brombal: To run Fourier Transform computation go to Data -> Data Analysis -> Fourier Analysis If Data Analysis tab is not present you should install the add in "Data Analysis"
Note on J4: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Disregard the first value (corresponding to frequency 0)

## ESF to LSF calculation
Date: 14/02/2023 | INPUT |  | OUTPUT
Author: Luca Brombal
Pixel No | ESF | Pixel No | LSF
0 | 0.010307 | 0 | 0
1 | 0.010236 | 1 | -7.1583e-05
2 | 0.0096059 | 2 | -0.00062996
3 | 0.010054 | 3 | 0.00044786
4 | 0.010127 | 4 | 7.3194e-05
5 | 0.010853 | 5 | 0.00072563
6 | 0.010329 | 6 | -0.00052324
7 | 0.010261 | 7 | -6.8588e-05
8 | 0.010523 | 8 | 0.00026191
9 | 0.01019 | 9 | -0.00033309
10 | 0.01058 | 10 | 0.00039067
11 | 0.011021 | 11 | 0.00044047
12 | 0.010012 | 12 | -0.0010085
13 | 0.010872 | 13 | 0.00086031
14 | 0.011369 | 14 | 0.00049689
15 | 0.010863 | 15 | -0.0005061
16 | 0.011328 | 16 | 0.00046484
17 | 0.010971 | 17 | -0.00035678
18 | 0.010972 | 18 | 9.4809e-07
19 | 0.010565 | 19 | -0.00040691
20 | 0.010637 | 20 | 7.1603e-05
21 | 0.011349 | 21 | 0.00071227
22 | 0.011843 | 22 | 0.00049424
23 | 0.010701 | 23 | -0.0011427
24 | 0.01142 | 24 | 0.00071924
25 | 0.011276 | 25 | -0.00014438
26 | 0.011 | 26 | -0.00027556
27 | 0.011293 | 27 | 0.00029271
28 | 0.011564 | 28 | 0.00027109
29 | 0.010859 | 29 | -0.00070469
30 | 0.010731 | 30 | -0.00012813
31 | 0.011376 | 31 | 0.00064464
32 | 0.010593 | 32 | -0.00078271
33 | 0.010784 | 33 | 0.00019075
34 | 0.011255 | 34 | 0.00047149
35 | 0.012118 | 35 | 0.00086263
36 | 0.011484 | 36 | -0.00063342
37 | 0.011329 | 37 | -0.00015577
38 | 0.011713 | 38 | 0.00038392
39 | 0.011815 | 39 | 0.00010201
40 | 0.011278 | 40 | -0.00053624
41 | 0.01069 | 41 | -0.0005883
42 | 0.011573 | 42 | 0.00088275
43 | 0.012082 | 43 | 0.00050872
44 | 0.01159 | 44 | -0.00049181
45 | 0.011509 | 45 | -8.067e-05
46 | 0.011562 | 46 | 5.3295e-05
47 | 0.011739 | 47 | 0.00017612
48 | 0.011582 | 48 | -0.00015682
49 | 0.012037 | 49 | 0.00045568
50 | 0.011929 | 50 | -0.00010834
51 | 0.011785 | 51 | -0.00014422
52 | 0.011216 | 52 | -0.00056851
53 | 0.013463 | 53 | 0.0022465
54 | 0.013024 | 54 | -0.00043885
55 | 0.012533 | 55 | -0.00049084
56 | 0.013245 | 56 | 0.0007118
57 | 0.012946 | 57 | -0.00029933
58 | 0.01207 | 58 | -0.00087608
59 | 0.012199 | 59 | 0.00012943
60 | 0.012533 | 60 | 0.00033429
61 | 0.013035 | 61 | 0.00050198
62 | 0.013601 | 62 | 0.00056541
63 | 0.012844 | 63 | -0.0007567
64 | 0.013786 | 64 | 0.0009418
65 | 0.013609 | 65 | -0.00017716
66 | 0.01476 | 66 | 0.0011509
67 | 0.013779 | 67 | -0.00098081
68 | 0.013283 | 68 | -0.000496
69 | 0.013871 | 69 | 0.00058831
70 | 0.013864 | 70 | -6.889e-06
71 | 0.014706 | 71 | 0.00084159
72 | 0.014812 | 72 | 0.00010577
73 | 0.014272 | 73 | -0.00053965
74 | 0.014731 | 74 | 0.00045933
75 | 0.013609 | 75 | -0.0011225
76 | 0.014846 | 76 | 0.001237
77 | 0.014042 | 77 | -0.00080321
78 | 0.014779 | 78 | 0.00073699
79 | 0.014614 | 79 | -0.00016548
80 | 0.014289 | 80 | -0.00032477
81 | 0.01534 | 81 | 0.0010507
82 | 0.014975 | 82 | -0.00036492
83 | 0.015132 | 83 | 0.00015701
84 | 0.015028 | 84 | -0.00010423
85 | 0.015138 | 85 | 0.00011027
86 | 0.016172 | 86 | 0.001034
87 | 0.016707 | 87 | 0.00053486
88 | 0.016976 | 88 | 0.0002695
89 | 0.017068 | 89 | 9.1227e-05
90 | 0.01686 | 90 | -0.00020726
91 | 0.016651 | 91 | -0.00020952
92 | 0.01683 | 92 | 0.0001792
93 | 0.017384 | 93 | 0.00055363
94 | 0.017912 | 94 | 0.00052841
95 | 0.017879 | 95 | -3.2729e-05
96 | 0.017962 | 96 | 8.2515e-05
97 | 0.017658 | 97 | -0.00030422
98 | 0.018697 | 98 | 0.001039
99 | 0.019833 | 99 | 0.0011359
100 | 0.019737 | 100 | -9.5898e-05
101 | 0.020061 | 101 | 0.00032387
102 | 0.02141 | 102 | 0.0013493
103 | 0.021825 | 103 | 0.00041517
104 | 0.021461 | 104 | -0.00036413
105 | 0.022103 | 105 | 0.00064193
106 | 0.023501 | 106 | 0.0013984
107 | 0.023628 | 107 | 0.00012676
108 | 0.023764 | 108 | 0.00013569
109 | 0.024547 | 109 | 0.00078329
110 | 0.02605 | 110 | 0.0015033
111 | 0.026843 | 111 | 0.00079303
112 | 0.028229 | 112 | 0.0013856
113 | 0.027378 | 113 | -0.00085115
114 | 0.028618 | 114 | 0.0012398
115 | 0.030439 | 115 | 0.0018217
116 | 0.030432 | 116 | -7.5735e-06
117 | 0.031615 | 117 | 0.0011838
118 | 0.032817 | 118 | 0.0012015
119 | 0.035153 | 119 | 0.002336
120 | 0.036907 | 120 | 0.0017542
121 | 0.037989 | 121 | 0.0010822
122 | 0.040604 | 122 | 0.0026145
123 | 0.043515 | 123 | 0.0029112
124 | 0.046678 | 124 | 0.0031632
125 | 0.049511 | 125 | 0.0028333
126 | 0.054824 | 126 | 0.0053121
127 | 0.059952 | 127 | 0.0051282
128 | 0.064749 | 128 | 0.0047971
129 | 0.071133 | 129 | 0.006384
130 | 0.07746 | 130 | 0.006327
131 | 0.08384 | 131 | 0.0063802
132 | 0.090811 | 132 | 0.0069704
133 | 0.098654 | 133 | 0.0078436
134 | 0.10851 | 134 | 0.0098536
135 | 0.12192 | 135 | 0.013412
136 | 0.13784 | 136 | 0.015918
137 | 0.1584 | 137 | 0.020558
138 | 0.1867 | 138 | 0.028303
139 | 0.23476 | 139 | 0.048061
140 | 0.33428 | 140 | 0.099524
141 | 0.45623 | 141 | 0.12194
142 | 0.51659 | 142 | 0.060366
143 | 0.54938 | 143 | 0.03279
144 | 0.57219 | 144 | 0.022808
145 | 0.58938 | 145 | 0.017193
146 | 0.60246 | 146 | 0.013076
147 | 0.61365 | 147 | 0.011186
148 | 0.6212 | 148 | 0.0075542
149 | 0.63163 | 149 | 0.010433
150 | 0.63681 | 150 | 0.0051782
151 | 0.638 | 151 | 0.0011919
152 | 0.64204 | 152 | 0.0040349
153 | 0.6466 | 153 | 0.0045647
154 | 0.65258 | 154 | 0.0059729
155 | 0.65593 | 155 | 0.0033597
156 | 0.65986 | 156 | 0.0039217
157 | 0.662 | 157 | 0.0021422
158 | 0.665 | 158 | 0.003002
159 | 0.6666 | 159 | 0.001599
160 | 0.66877 | 160 | 0.0021704
161 | 0.66648 | 161 | -0.0022857
162 | 0.66463 | 162 | -0.0018575
163 | 0.66689 | 163 | 0.0022665
164 | 0.66512 | 164 | -0.0017756
165 | 0.66855 | 165 | 0.0034271
166 | 0.67217 | 166 | 0.0036274
167 | 0.67521 | 167 | 0.003039
168 | 0.67757 | 168 | 0.0023565
169 | 0.67792 | 169 | 0.00035578
170 | 0.67916 | 170 | 0.0012374
171 | 0.68052 | 171 | 0.0013582
172 | 0.68196 | 172 | 0.0014448
173 | 0.6856 | 173 | 0.0036312
174 | 0.6853 | 174 | -0.00029147
175 | 0.68949 | 175 | 0.0041889
176 | 0.69001 | 176 | 0.00051689
177 | 0.6907 | 177 | 0.00069165
178 | 0.69038 | 178 | -0.00032115
179 | 0.68587 | 179 | -0.0045134
180 | 0.68838 | 180 | 0.0025142
181 | 0.68858 | 181 | 0.00019789
182 | 0.68924 | 182 | 0.00066274
183 | 0.69257 | 183 | 0.0033253
184 | 0.69593 | 184 | 0.0033608
185 | 0.69598 | 185 | 5.4657e-05
186 | 0.69695 | 186 | 0.00096285
187 | 0.69434 | 187 | -0.0026011
188 | 0.6967 | 188 | 0.0023586
189 | 0.69631 | 189 | -0.00038922
190 | 0.69564 | 190 | -0.00067824
191 | 0.69528 | 191 | -0.0003534
192 | 0.69967 | 192 | 0.0043834
193 | 0.69736 | 193 | -0.0023066
194 | 0.69872 | 194 | 0.0013642
195 | 0.69856 | 195 | -0.00016326
196 | 0.69508 | 196 | -0.0034759
197 | 0.69415 | 197 | -0.00093371
198 | 0.69341 | 198 | -0.00074196
199 | 0.6936 | 199 | 0.00019634
200 | 0.69578 | 200 | 0.0021781
201 | 0.69631 | 201 | 0.0005306
202 | 0.69604 | 202 | -0.00027096
203 | 0.69689 | 203 | 0.00085068
204 | 0.69459 | 204 | -0.0023017
205 | 0.69193 | 205 | -0.0026624
206 | 0.69361 | 206 | 0.0016838
207 | 0.69264 | 207 | -0.00097513
208 | 0.69378 | 208 | 0.0011405
209 | 0.69378 | 209 | 6.0797e-06
210 | 0.69578 | 210 | 0.0019961
211 | 0.6947 | 211 | -0.0010846
212 | 0.69474 | 212 | 4.8816e-05
213 | 0.69619 | 213 | 0.0014499
214 | 0.69907 | 214 | 0.002875
215 | 0.69936 | 215 | 0.00029469
216 | 0.69982 | 216 | 0.00046098
217 | 0.70104 | 217 | 0.0012168
218 | 0.69925 | 218 | -0.0017959
219 | 0.70035 | 219 | 0.0011011
220 | 0.70286 | 220 | 0.0025125
221 | 0.70335 | 221 | 0.00049525
222 | 0.70058 | 222 | -0.002773
223 | 0.69744 | 223 | -0.0031417
224 | 0.69563 | 224 | -0.0018121
225 | 0.69673 | 225 | 0.0011027
226 | 0.69682 | 226 | 8.7023e-05
227 | 0.69847 | 227 | 0.0016575
228 | 0.69914 | 228 | 0.00066954
229 | 0.69812 | 229 | -0.0010286
230 | 0.69711 | 230 | -0.0010079
231 | 0.69898 | 231 | 0.0018722
232 | 0.69929 | 232 | 0.00031453
233 | 0.69783 | 233 | -0.001469
234 | 0.69773 | 234 | -9.8109e-05
235 | 0.69656 | 235 | -0.0011725
236 | 0.6971 | 236 | 0.00054139
237 | 0.69774 | 237 | 0.00064778
238 | 0.69926 | 238 | 0.0015155
239 | 0.69918 | 239 | -8.4579e-05
240 | 0.69937 | 240 | 0.0001933
241 | 0.69994 | 241 | 0.00057071
242 | 0.69868 | 242 | -0.0012555
243 | 0.69797 | 243 | -0.00071436
244 | 0.70022 | 244 | 0.0022515
245 | 0.7018 | 245 | 0.0015814
246 | 0.70011 | 246 | -0.0016907
247 | 0.702 | 247 | 0.0018917
248 | 0.69986 | 248 | -0.0021461
249 | 0.70028 | 249 | 0.00042349
250 | 0.70035 | 250 | 6.4552e-05
251 | 0.69687 | 251 | -0.0034784
252 | 0.6953 | 252 | -0.0015648
253 | 0.69694 | 253 | 0.0016376
254 | 0.69802 | 254 | 0.0010824
255 | 0.70228 | 255 | 0.0042539
256 | 0.6984 | 256
257 | 0.70032 | 257
258 | 0.69951 | 258
259 | 0.70155 | 259
260 | 0.701 | 260
261 | 0.69638 | 261
262 | 0.69744 | 262
263 | 0.69892 | 263
264 | 0.70159 | 264
265 | 0.70243 | 265
266 | 0.69817 | 266
267 | 0.69696 | 267
268 | 0.69737 | 268
269 | 0.69858 | 269
270 | 0.70071 | 270
271 | 0.70105 | 271
272 | 0.70008 | 272
273 | 0.69837 | 273
274 | 0.70104 | 274
275 | 0.70297 | 275
276 | 0.7034 | 276
277 | 0.70291 | 277
278 | 0.70069 | 278
279 | 0.70064 | 279
280 | 0.70103 | 280
281 | 0.70363 | 281